Instruction Manual for Pulsatile Flow

Version 1.5.0

1   Working principle and installation

As described  in the general  instruction manual the  ibidi Pump and  the ibidi Fluidic  Unit are

designed  to work  together  to create  a  flow –  unidirectional,  oscillating,  or pulsating  –  of

medium within  the channel  of the ibidi  slides. In  order to create  a pulsatile  flow you need

one pump and two  Fluidic Units. The first Fluidic  Unit will create a unidirectional flow  which

the second Fluidic Unit will periodically interrupt to create pulsation.

Please follow these steps to set up a pulsatile flow experiment:

1.   Connect the pump to your  computer and the two Fluidic  Units to the pump. Plug for

example the first Fluidic Unit to ‘Port 1’ of the pump and the second to ‘Port 2’.

2.   Connect the air pressure tubing only to the first Fluidic Unit.

3.   Mount the Perfusion Set on the first Fluidic Unit.

•

This  experiment  works  best  with  long  Perfusion  Sets  like  ‘yellow/green’,

‘white’, and ‘black’.

4.   Using the pump and  the first Fluidic Unit set up  a continuous and unidirectional flow

with  the  desired  flow  rate  and  the  resulting  shear  stress  with  the  help  of   the

PumpControl software.

5.   Route the  tubing of  the Perfusion  Set which  is coming  out of  the first  Fluidic Unit

through the pinch valve of the second Fluidic  Unit. Make sure that you use either the

two front or  the two rear  openings of the pinch  valve for that. Do  not intermix front

and rear openings. Compare your set-up with Figure 1.

6.   Run the second Fluidic Unit in the PumpControl software  as an ‘oscillating’ Unit. This

way  you can  select the  switching time  in  the range  of 0.2s  –  5s which  results in

pulsation periods  of 0.4s  – 10s. This  in turn  corresponds to  frequencies of 2.5Hz  –

0.1Hz. Set settings are shown in Figure 2.

ibidi instruction manual  – pulsatile flow

© ibidi GmbH 2012, version 1.5.0, revision 0, 2012-04-19, Dr. Miro Kroutvar

1

Figure 1 Set-up for pulsatile flow. Continuous flow in the Perfusion Set

coming out of Fluidic Unit 1 is periodically pinched off by Fluidic Unit 2.

ibidi instruction manual – pulsatile flow

ibidi GmbH 2012, version 1.5.0, revision 0, 2012-04-19, Dr. Miro Kroutvar

2

Fluidic Unit 1 creates a

continuous flow

Fluidic Unit 2 creates

pulsation (here: 1.0Hz)

Figure 2 Settings for PumpControl

ibidi instruction manual  – pulsatile flow

© ibidi GmbH 2012, version 1.5.0, revision 0, 2012-04-19, Dr. Miro Kroutvar

3