Instructions

3WellChamber,removable

The3WellChamber,removableisaremovablesiliconechamber,mountedonaglassslide,for

cellcultureandimmunoﬂuorescencestainings.Itallowstheuseofstandardcultivation,staining

andmountingtechniqueswithcoverslipsealing.

Aftermountingtheglassslidewithacoverslip,itisreadyfortheusewithuprightandinverted

microscopes,aswellasforlongtermstorage.Suitable24mm×60mmcoverslipsareprovidedby

ibidi(10811).

Thespecialgeometryofthe3WellChamber,removableprovidesastructureinsideeachwell,formingaroundinnerwell.

Theoptionofplacingaroundcoverslip(ibidi,     10815)intothewellallowstheusertoﬁlltheinnerwellexclusively,       by

minimizingthevolumeandeliminatingthemeniscus.

Material

DimensionsoftheSingleWells

Dimensionofwells(w×l×h)

Volumeperwell

16.5×16.5×8mm

1100µl

The 3 Well Chamber, removable is comprised of a self–

adhesive silicone gasket mounted on a standard mi-

croscopyglassslide.Thegasketismanufacturedfrombio-

compatiblesiliconematerial. Althoughbothmaterialsare

autoclavableandcompatiblewithalcohols,wedonotrec-

ommendreusingthem. Theglassbottomexhibitsground

edgesandtwinfrostedends.

Reducedvolumeperwell

Growthareaperwell

Coatingareaperwell

Diameterofinnerwell

Heightoftheinnerwell

150µl

1.66cm2

3.37cm2

13.5mm

0.75mm

Geometry

ShippingandStorage

The3WellChamber,removableprovidesastandardslide

formataccordingtoISO8037/1.

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

GeneralDimensions

Numberofwells

3

Conditions

Bottomsize(w×l×h)

Twinfrostedends

Totalheightwithlid

26×76×1mm

13mm

Shippingconditions

Storageconditions

Ambient

11mm

RT(15-25°C)

Thesinglewellsofthe3WellChamber,removableconsist

oftheouterrectangularwellwithaninletformingaround

innerwell. Whenplacingacoverslipwith15mmdiame-

terontotheinlet,  aseparatechamberislockedup.    This

assuresahomogeneousdistributionofcellsorsubstances

andminimizestheliquidvolume.   Theseparatechamber

canbeaccessedviafourinletsintheedges.

ShelfLifeofDifferentSurfaces

ibiTreat,GlassBottom,ESS

36months

18months

Collagen,Poly-L-Lysine

SurfaceandCoating

The 3 Well Chamber, removable is mounted on an un-

coatedglassslidewithtwinfrostedends.    Washingsteps

(e.g. withPBS)beforecellseedingcanremoveglassdust

whichisadvantageousfordirectcellgrowthonthesur-

face.

Protein coatings increase direct cell growth of adherent

cells.Speciﬁccoatingsonglassarepossiblefollowingthis

protocol:

Left:Singlewellﬁlledwith1.1mlmedium.Right:Volume

reductionbyuseofthecoverslip(150µlmedium).

3WellChamber,removable

Page1

Version1.0(2016-02-02)

Instructions

3WellChamber,removable

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.  Prepareyour

slide. Adjusttheconcentrationtoacoatingareaof

3.37cm2andacoatingvolumeof1100µlperwell.

• GrabthecoverslipwiththePick-UpToolandplace

itontothewellinlet.  ThePick-UpToolworkslikea

Pasteurpipette.

• Injectavolumeof150µlintooneoftheedges,that

arekeptopen.

• Apply1100µlintoeachwell.Makesurethattheen-

tirebottomiscoveredwithliquidbyslightlytilting

orshakingtheslide.Putonthelidandleavetheslide

atroomtemperatureforatleast30minutes.

• Aspirate the solution and wash with the recom-

mendedproteindilutionbuffer.   Optionally, letdry

atroomtemperature.  Attention, somecoatingpro-

teinsmightdegeneratewhendrying!

SeedingCells

• Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.Depending

onyourcelltype,applicationofa2-6×104cells/ml

suspensionshouldresultinaconﬂuentlayerwithin

2–3days.

Removethecoverslipasfollows:

• Add950µlinonecornernexttothecoverslip.

• GrabtheﬂoatingcoverslipwiththePick-UpToolor

tweezersandremoveitfromthewell.

• Apply1100µlcellsuspensionintoeachwellofthe

slide.Avoidshakingasthiswillresultininhomoge-

neousdistributionofthecells.

• Cover the slide with the supplied lid. Incubate at

37°Cand5%CO2asusual.

Undemanding cells can be left in their seeding medium

foruptothreedaysandgrowtoconﬂuencethere.     How-

ever, best results might be achieved when the medium

is changed every 1–2 days. Carefully aspirate the old

mediumandreplaceitby1100µlfreshmediumperwell.

Note:

SolventsforFixation,StainingandOtherPur-

poses

For a more homogeneous cell distribution use the

15mm coverslips (ibidi, 10815). Autoclave the cov-

erslipspriortouse.  Pleasefollowtheprotocolbelow.

Use150µlofcellsuspensionwithaconcentrationof

1.7-3.8×105cells/ml.

3 Well Chamber, removable is compatible to methanol,

acetone,acids,alkalis,PFA,DMSO,siliconeoil,andmin-

eraloilforcellculture.

ImmunoﬂuorescenceMicroscopy

HandlingoftheRoundCoverslip

Aftercultivation,thecellscanbeﬁxedandstainedintwo

differentways.

Forﬁllingthechamberbyuseoftheroundcoverslip,fol-

lowthisprotocol.OneCoverslipPick-UpToolisdelivered

witheveryunitof15mmcoverslips(ibidi,10815):

Singlewelltechnique:

Allstepsarecarriedoutinthesinglewellsbeforeremov-

ing the silicone gasket. Using the 15mm coverslips will

reducethevolumeofthestainingsolutionto150µl.

• Beforeapplyingthecoverslip,thewellmustbeemp-

tied.

3WellChamber,removable

Page2

Version1.0(2016-02-02)

Instructions

3WellChamber,removable

• Perform your standard staining protocol (ﬁxation,

permeabilization, staining, washing) in the single

wells(1).

• Starting from one edge, remove the silicone gasket

byhandorbyusingtweezers(2).

(1)

(3)

Mounting:

Mounttheslidewithacoverslip24mm×60mm(ibidi,

10811)andapermanentmountingmediumofyourchoice

(4).

(2)

Note:

Ifyouwanttoreducethevolumeofthestainingsolu-

tionneeded,usetheroundcoverslipasshowninthe

coversliphandlingprotocol.Washthewellbyadding

950µlandremovingthecoverslip.

(4)

Paralleltechnique:

Tip:

All necessary steps (ﬁxation, permeabilization, staining,

washing)arecarriedoutbydippingthewholeslideinto

thesolutionsafterremovingthesiliconegasket(4).

For mounting of slide samples, a hardening per-

manent mounting medium such as Fluoroshield™

(Sigma-Aldrich), Vectashield® (Vector Laboratories

Inc.)orProLong®Antifade(ThermoFisherScientiﬁc)

isrecommended.

• Starting from one edge, remove the silicone gasket

byhandorbyusingtweezers(2).

ibidiMountingMediumisnotrecommendedbecause

itisnon–hardeningandstaysaliquid(whichisad-

vantageousforµ–Slidesandµ–Dishes).

• Perform your standard staining protocol (ﬁxation,

permeabilization,staining,washing)bydippingthe

slideintothesolutions(3).

3WellChamber,removable

Page3

Version1.0(2016-02-02)

Instructions

3WellChamber,removable

OrderingInformation

Cat.No.

Description

80381

3WellChamber,removable:microscopyglassslide,sterilized

80841

81201

10811

10815

8WellChamber,removable:microscopyglassslide,sterilized

12WellChamber,removable:microscopyglassslide,sterilized

Coverslips60mm,unsterileforChambers,removable,#1.5H(170±5µm)D263MSchottglass,24mm×

Coverslips and Coverslip Pick-Up Tool for 3Well Chamber, removable: # 1.5H (170  ±

5µm)D263MSchottglass,15mm,unsterile,includingonesiliconerubbertoolforconve-

nientglasscoversliphandling

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

3WellChamber,removable

Page4

Version1.0(2016-02-02)