Instructions

µ–Slide2Well  Ph+

Theibidiproductfamilyiscomprisedofavarietyofµ–Slidesandµ–Dishes,whichhave

all been designed for high–end microscopic analysis of ﬁxed or living cells. The high

opticalqualityofthematerialissimilartothatofglass,soyoucanperformallkindsof

ﬂuorescenceexperimentswithuncompromisedresolutionandchoiceofwavelength.

Theµ–Slide2WellPh+(Phasecontrastplus)isanarrayof2squareﬁeldswherecellscanbe

cultivatedandinvestigatedwithmicroscopicalmethods.Theµ–Slide2WellPh+improves

theopticalqualityofphasecontrastmicroscopy.   Incontrasttotheclassicµ–Slide2Well,

thePh+versionprovidesaspecialplateinthecenterofthewells.      Thisplatesuppresses

themeniscuswhichisdisturbingthephasecontrasteffectinnormalopenwells.Openingsnearthecornersprovideaccess

tothewellsforﬁllingandaspiratingliquidseasily.

Material

ShelfLifeofDifferentSurfaces

ibiTreat,GlassBottom,ESS

Collagen,Poly-Lysine

Fibronectin

36months

18months

4months

ibidiµ–Slides,µ–Dishes,andµ–Platesaremadeofaplas-

ticthathasthehighestopticalquality.    Thepolymercov-

ersliponthebottomexhibitsextremelylowbirefringence

and autoﬂuorescence, similar to that of glass. Also, it is

not possible to detach the bottom from the upper part.

Theµ–Slides,µ–Dishes,andµ–Platesarenotautoclavable,

sincetheyareonlytemperature–stableupto80°C/175°F.

Please note that gas exchange between the medium and

incubator’satmosphereoccurspartiallythroughthepoly-

mercoverslip,whichshouldnotbecovered.

Geometry

Theµ–Slide2WellPh+providesastandardslideformatac-

cordingtoISO8037/1.

Geometryofµ–Slide2WellPh+

OpticalPropertiesibidiPolymerCoverslip

Numberofwells

2

RefractiveindexnD(589nm)

Abbenumber

1.52

Dimensionsofwells(w×l×h)inmm     21.2 × 23.3 ×

3.0

56

Thickness

No.1.5(180µm)

polymercoverslip

Growthareaperwell

Coatingareaperwell

Volumeperwell

4.8cm2

11.4cm2

Material

1.5ml

Pleasenote! Theibidipolymercoverslipiscompatible

withcertaintypesofimmersionoilonly.    Alistofsuit-

ableoilscanbefoundonpage3.

Liquidheight

3.0mm

10.8mm

No.1.5

Totalheightwithlid

Bottommatchescoverslip

µ–SlideSurfaces

ShippingandStorage

Dependingonthetypeofcellsandthespecialapplication

youareusing,youwillneedµ–Slideswithdifferentsur-

faces.Ifyoudonotrequireanyspecialadhesionmolecules

foryourapplication,thebestchoicewillbeibiTreat,atis-

sueculturetreatedsurface.

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

Weprovideprecoatedµ–Slideswithadhesionsubstrates

likeCollagenIV,Fibronectin,Poly–L–Lysin,andPoly–D–

Lysin.Suchadhesionsubstrateshavebeenshowntostim-

ulatetheadhesionandgrowthofvariouscelllinesinµ–

Slides.Onlyhigh–qualitysubstratesareused1.

Conditions

Shippingconditions

Storageconditions

Ambient

RT(15-25°C)

1CollagenIV:Corning#356233,Fibronectin:Corning#354008,Poly–L–Lysin:Sigma#P4832,Poly–D–Lysin:Corning#354210

µ–Slide2WellPh+

Page1

Version2.0(2015-11-22)

Instructions

µ–Slide2Well  Ph+

Theuncoatedµ–Slideismanufacturedfromhydrophobic

plastic. For the cultivation of most cell lines, it is indis-

pensabletotreattheuncoatedµ–Slidewithbiopolymers,

whichmediatecelladhesionandgrowth.

Tip:

Thedaybeforeseedingthecellswerecommendplac-

ingthecellmediumandtheµ–Slideintotheincubator

forequilibration. Thiswillpreventtheliquidinside

fromemergingairbubblesovertheincubationtime.

Coatingyourµ–Slide2Well    Ph+

Theuncoatedµ–Slidemustbecoatedtopromotecellad-

hesion.Ifyouwanttoestablishacertaincoatingtomatch

your needs, we recommend testing your coating proce-

dure on both uncoated and ibiTreat µ–Slides, since we

haveobservedthatsomebiomoleculesadheredifferently

tohydrophobicandhydrophilicplasticsurfaces.

FillingandHandling

Fill the wells by using a standard pipet. Inject the cell

suspensiondirectlyintooneoftheopenings.Mediumex-

changeiseasilydonebyaspiratingtheentirevolumeand

reﬁllingusing1.5mlperwell.

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.

• Apply1.5mlperwellandleaveatroomtemperature

foratleast30minutes.

• Aspirate the solution and wash with the recom-

mendedproteindilutionbuffer.

• Optionallyletdryatroomtemperature.    Attention,

somecoatingproteinsmightdegeneratewhendry-

ing!

DetailedinformationaboutcoatingsisprovidedinAppli-

cationNote08

SeedingCells

• Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.Depending

onyourcelltype,applicationofa5-11×104cells/ml

suspensionshouldresultinaconﬂuentlayerwithin

2–3days.

• Apply1.5mlcellsuspensionintoeachwelloftheµ–

Slide.Avoidshakingasthiswillresultininhomoge-

neousdistributionofthecells.

• Cover the slide with the supplied lid. Incubate at

37°Cand5%CO2asusual.

Undemanding cells can be left in their seeding medium

foruptothreedaysandgrowtoconﬂuencethere.     How-

ever, best results might be achieved when the medium

is changed every 1–2 days. Carefully aspirate the old

mediumandreplaceitby1.5ml/wellfreshmedium.

µ–Slide2WellPh+

Page2

Version2.0(2015-11-22)

Instructions

µ–Slide2Well  Ph+

PreparationforCellMicroscopy

µ–Slide2well Ph+selectionguide

µ–Slide2Well

µ–Slide2WellPh+

To analyze your cells, no special preparations are neces-

sary.Cellscanbeobservedlive,orﬁxeddirectlyintheµ–

Slideonaninvertedmicroscope.  Youcanuseanyﬁxative

of your choice. The µ–Slide material is compatible with

avarietyofchemicals,e.g.,acetoneormethanol.    Further

speciﬁcationscanbefoundatDuetothe

thinbottomofonly180µm,highresolutionmicroscopyis

possible.

Standard open wells for maxi-  Specialplateinthecenterofthe

mum sample access.

Meniscus

wells suppresses meniscus for-

mation. Nomeniscus–parallel

disturbs the beam path.

Good

phasecontrastqualityonlyinthe    beampath.  Forexcellentphase

centerofeachwell

contrastmicroscopyalloverthe

wells.

ImmersionOil

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

Company

Product

OrderingNumber

Zeiss

Zeiss

Leica

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

µ–Slide2WellPh+

Page3

Version2.0(2015-11-22)

Instructions

µ–Slide2Well  Ph+

OrderingInformation

Theµ–Slide2WellisavailableasopenwellandasaPh+version,aswellasinaglassbottomversion.          Seetablebelowfor

choosingyourµ–Slide2Well.

µ–Slide2Well

Cat.No.

Description

80886

80882

80883

80884

80885

80881

80887

µ–Slide2WellibiTreat:#1.5polymercoverslip,tissueculturetreated,sterilized

µ–Slide2WellCollagenIV:#1.5polymercoverslip,sterilized

µ–Slide2WellFibronectin:#1.5polymercoverslip,sterilized∗

µ–Slide2WellPoly-L-Lysine:#1.5polymercoverslip,sterilized

µ–Slide2WellPoly-D-Lysine:#1.5polymercoverslip,sterilized∗

µ–Slide2WellUncoated:#1.5polymercoverslip,hydrophobic,sterilized

µ–Slide2WellGlassBottom:1.5H(170µm±5µm)D263MSchottglass,sterilized

∗availableonrequestonly

µ–Slide2WellPh+

Cat.No.

Description

80296

80292

80293

80294

80295

80291

80297

µ–Slide2WellPh+ibiTreat:#1.5polymercoverslip,tissueculturetreated,sterilized

µ–Slide2WellPh+CollagenIV:#1.5polymercoverslip,sterilized

µ–Slide2WellPh+Fibronectin:#1.5polymercoverslip,sterilized∗

µ–Slide2WellPh+Poly-L-Lysine:#1.5polymercoverslip,sterilized

µ–Slide2WellPh+Poly-D-Lysine:#1.5polymercoverslip,sterilized∗

µ–Slide2WellPh+Uncoated:#1.5polymercoverslip,hydrophobic,sterilized

µ–Slide2WellPh+GlassBottom:1.5H(170µm±5µm)D263MSchottglass,sterilized

∗availableonrequestonly

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

µ–Slide2WellPh+

Page4

Version2.0(2015-11-22)