Instructions

µ–Slide18Well-Flat

Theibidiproductfamilyiscomprisedofavarietyofµ–Slidesandµ–Dishes,whichhaveall

been designed for high–end microscopic analysis of ﬁxed or living cells. The high optical

qualityofthematerialissimilartothatofglass,soyoucanperformallkindsofﬂuorescence

experimentswithuncompromisedresolutionandchoiceofwavelength.Theµ–Slide18Well-

Flatisaconvenientmulti–wellplateforquickandeasyteststoprobeexperimentalparameters

likeantibodydilution,  seedingdensityormosteffectivedrugconcentrations.    Unlikein96

wellplates,thewholeareaofthewellcanbeobservedinphase–contrastmicroscopymode

sincethereisnodisturbingrefractionoftheliquidsurfacemeniscus.

Material

ShelfLifeofDifferentSurfaces

ibiTreat,GlassBottom,ESS

Collagen,Poly-Lysine

Fibronectin

36months

18months

4months

ibidiµ–Slides,µ–Dishes,andµ–Platesaremadeofaplas-

tic that has the highest optical quality. The bottom ma-

terial exhibits extremely low birefringence and autoﬂuo-

rescence, similar to that of glass. Also, it is not possible

todetachthebottomfromtheupperpart.     Theµ–Slides,

µ–Dishes, and µ–Plates are not autoclavable, since they

areonlytemperature–stableupto80°C/175°F.Pleasenote

that gas exchange between the medium and incubator’s

atmosphere occurs partially through the polymer cover-

slip,whichshouldnotbecovered.

Geometryoftheµ–Slide18Well-Flat

Theµ–Slide18Well-Flatprovidesstandardslideformat

accordingtoISO8037/1.Thewelltowelldistanceof9mm

(like96wellplates)allowsusingmultichannelpipettes.

OpticalPropertiesibidiStandardBottom

Geometryoftheµ–Slide18Well-Flat

RefractiveindexnD(589nm)

Abbenumber

1.52

Numberofwells

18

56

Volumeperwell

30µl

Thickness

No.1.5(180µm)

Welldiameter

5mm

Material

microscopyplastic/

polymercoverslip

0.2cm2

0.25cm2

Growthareaperwell

Coatingareaperwell

Heightwith/withoutlid

Bottommatchescoverslip

5.0/1.6mm

No.1.5

Please note! The ibidi Standard Bottom is compatible

withcertaintypesofimmersionoilonly.    Alistofsuit-

ableoilscanbefoundonpage2.

µ–SlideSurfaces

Dependingonthetypeofcellsandthespecialapplication

youareusing,youwillneedµ–Slideswithdifferentsur-

faces.Ifyoudonotrequireanyspecialadhesionmolecules

foryourapplication,thebestchoicewillbeibiTreat,atis-

sueculturetreatedsurface.

ShippingandStorage

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

Weprovideprecoatedµ–Slideswithadhesionsubstrates

likeCollagenIV,Fibronectin,Poly–L–Lysin,andPoly–D–

Lysin.Suchadhesionsubstrateshavebeenshowntostim-

ulatetheadhesionandgrowthofvariouscelllinesinµ–

Slides.Onlyhigh–qualitysubstratesareused1.

Conditions

Shippingconditions

Storageconditions

Ambient

Theuncoatedµ–Slideismanufacturedfromhydrophobic

plastic. For the cultivation of most cell lines, it is indis-

RT(15-25°C)

1CollagenIV:Corning#356233,Fibronectin:Corning#354008,Poly–L–Lysin:Sigma#P4832,Poly–D–Lysin:Corning#354210

µ–Slide18Well-Flat

Page1

Version2015-08-05

Instructions

µ–Slide18Well-Flat

pensabletotreattheuncoatedµ–Slidewithbiopolymers,

whichmediatecelladhesionandgrowth.

is changed every 1–2 days. Carefully aspirate the old

mediumandreplaceitby30µl/wellfreshmedium.

Thevolumeofasinglewellisverysmall.     Dependingon

yourcelltypethemediummightbeconsumedaftersome

hours. Ifyouwanttoincubateyourcellsforlongerthan

acoupleofhourswerecommendtoaspirateandreﬁllcell

mediumeveryday.

Coatingyourµ–Slide18Well-Flat

Theuncoatedµ–Slidemustbecoatedtopromotecellad-

hesion.Ifyouwanttoestablishacertaincoatingtomatch

your needs, we recommend testing your coating proce-

dure on both uncoated and ibiTreat µ–Slides, since we

haveobservedthatsomebiomoleculesadheredifferently

tohydrophobicandhydrophilicplasticsurfaces.

Tip:

Ifthewellsareproperlyﬁlledwith30µl,     theliquid

surfaceisplanarandingoodalignmentwiththeµ–

Slidessurface.Thisishowyouwillbeabletoobserve

thewholewellareawithunimpairedphasecontrast.

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.

• Apply30µlperwellandleaveatroomtemperature

foratleast30minutes.

• Aspirate the solution and wash with the recom-

mendedproteindilutionbuffer.

PreparationforCellMicroscopy

• Optionallyletdryatroomtemperature.    Attention,

somecoatingproteinsmightdegeneratewhendry-

ing!

To analyze your cells, no special preparations are nec-

essary. Cells can be observed live, or ﬁxed directly in

the µ–Slide preferably on an inverted microscope. You

can use any ﬁxative of your choice. The µ–Slide mate-

rial is compatible with a variety of chemicals, e.g., ace-

toneormethanol.  Furtherspeciﬁcationscanbefoundat

Duetothethinbottomofonly180µm,

highresolutionmicroscopyispossible.

DetailedinformationaboutcoatingsisprovidedinAppli-

cationNote08

Seedingcells

• Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.Depending

onyourcelltype,applicationofa2-4×105cells/ml

ImmersionOil

suspensionshouldresultinaconﬂuentlayerwithin

2–3days.

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

• Apply30µlcellsuspensionintoeachwelloftheµ–

Slide.Avoidshakingasthiswillresultininhomoge-

neousdistributionofthecells.

Company

Product

OrderingNumber

• Cover the slide with the supplied lid. Incubate at

37°Cand5%CO2asusual.

Zeiss

Zeiss

Leica

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

Undemanding cells can be left in their seeding medium

foruptothreedaysandgrowtoconﬂuencethere.     How-

ever, best results might be achieved when the medium

µ–Slide18Well-Flat

Page2

Version2015-08-05

Instructions

µ–Slide18Well-Flat

OrderingInformation

Theµ–Slide18Well-Flatfamilyisavailablewithdifferentsurfaces.      Seetablebelowforchoosingyourµ–Slide18Well-

Flat.

Cat.No.

Description

81826

81822

81823

81824

81825

81821

81831

µ–Slide18Well-FlatibiTreat:#1.5polymercoverslip,tissueculturetreated,sterilized

µ–Slide18Well-FlatCollagenIV:#1.5polymercoverslip,sterilized

µ–Slide18Well-FlatFibronectin:#1.5polymercoverslip,sterilized∗

µ–Slide18Well-FlatPoly–L–Lysine:#1.5polymercoverslip,sterilized

µ–Slide18Well-FlatPoly–D–Lysine:#1.5polymercoverslip,sterilized∗

µ–Slide18Well-FlatUncoated:#1.5polymercoverslip,hydrophobic,sterilized

µ–Slide18Well-FlatMicrodissection:PEN-membrane,sterilized∗∗

∗availableonrequestonly

∗∗ThePENfoildoesnotﬁttostandardcoverslipthickness.

µ–Slide18Well-Flat

Page3

Version2015-08-05

Instructions

µ–Slide18Well-Flat

SelectedReferences

P.Carrara,P.Stano,andP.L.Luisi.     GiantVesicles”Colonies”: AModelforPrimitiveCellCommunities.    ChemBioChem,

2012.doi:10.1002/cbic.201200133.

S. Frey and D. Go¨rlich. A Saturated FG-Repeat Hydrogel Can Reproduce the Permeability Properties of Nuclear Pore

Complexes.Cell ,2007.

S.Sucic,S.Dallinger,B.Zdrazil,R.Weissensteiner,T.N.Jorgensen,M.Holy,O.Kudlacek,S.Seidel,J.H.Cha,andU.Gether.

Theaminoterminusofmonoaminetransportersisaleverrequiredfortheactionofamphetamines.       JournalofBiological

Chemistry,2010.doi:10.1074/jbc.M109.083154.

P.J.Tesar.Derivationofgerm-line-competentembryonicstemcelllinesfrompreblastocystmouseembryos.ProcNatlAcad

SciUSA ,2005.

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

µ–Slide18Well-Flat

Page4

Version2015-08-05