Application Note 14

2D Chemotaxis assay using μ-Slide Chemotaxis 2D

0.  Important Notes

•

•

•

•

Read the related document “Important Notes” first.

Follow all steps in this protocol carefully.

Watch the handling movie on www.ibidi.com.

Start with free samples and the included food coloring.

1.  General Information

The μ-Slide   Chemotaxis  2D  is  a  tool  for  observing

chemotactical responses of adherent migrating cells

in  2D  over   extended  periods  of  time.   The  linear

concentration     profile    which     is     required     for

chemotactical  movement  is  generated  by  diffusion

and stable for at least 48 hours. Three chambers on

one slide allow parallel chemotaxis experiments. The

following protocol is adapted for adherent cells only.

Color code for liquids in

this Application Note:

Chamber configuration:

1 Observation area (2000x1000μm

2 Large reservoir (40μl)

2

)

3 Large reservoir (40μl)

4 Filling port with small side channel

5 Filling port with small side channel

6 Filling port

7 Filling port

8 Plugs

9 Cultivation caps

2.  Principle

Two  large  volume  reservoirs  are  connected  by  a  thin  slit.  The  reservoirs  (2,  3)  contain

different chemoattractant concentrations (indicated below by red and blue color). Inside the

connecting slit a linear and stable concentration profile is generated by diffusion.

Cross  section   indicating  the   place  of

cells, chemoattractant, and gradient

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 1 of 15

Application Note 14

3.  Equipment

It is 100% necessary using the following equipment to conduct

chemotaxis experiments with the μ-Slide Chemotaxis 2D:

•

•

•

Adherent cells

Optimal conditions for cells (heated or incubated stage)

A  humid  chamber,  such  as  a  10  cm  Petri  dish  with  wet

tissue (as shown on the right)

•

•

Inverted microscope (phase contrast, fluorescence, … )

Time lapse video equipment (CCD camera, video camera,

acquisition software)

Humid Chamber

With wet sterile tissue.

•

Use only correct 10 - 200 μl pipet tips. Others do not work.

o    Greiner Bio-One 739261, 739280, 739290, 772288 or

related beveled Greiner tips.

o    Axygen T-200-C, TR-222-C, TR-222-Y or related Axygen

beveled tips.

o    STARLAB TipOne RPT S1161-1800 or related beveled

TipOne tips.

o    Sorenson  BioScience   MulTi  Fit   Tip  10590,   15320,

15330.

•

•

Slant cosmetics tweezers, for plug handling

Optional: Motorized stage and autofocus (x,y,z) to observe

all 3 chambers in parallel.

Beveled Pipet Tip

See list for correct models

Correct tip

¾  Fits on top

Wrong tip

¾  Fits too tightly

¾ Closes completely   ¾  Gets stuck

Slant Tweezers

For convenient plug handling

¾  Does not stuck

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 2 of 15

Application Note 14

4.  Coating μ-Slide Chemotaxis

For  direct  use  the  tissue  culture  treated  version μ-Slide

Chemotaxis 2D, ibiTreat (80306) or the precoated μ-Slide

Chemotaxis 2D, Collagen IV, sterile (80302) can be used.

Other   coatings   can  be   adapted   by   using   the   same

procedure described below. Work under sterile conditions.

a)   Full Coating Protocol with Collagen IV

Please use the following protocol to coat the entire inside

of the chamber with Collagen IV.

2

(μ-Slide   Chemotaxis   2D,  ibiTreat,   0.8  μg/cm ,   80  μl,

24 μg/ml, 2.4 cm

2

):

•

•

Dilute the Collagen IV (e.g. BD Cat.-Nr. 356233) to

a concentration of 24 μg/ml using 0.05 M HCl.

Inject 80 μl coating solution per chamber using the

filling   port   5   (step    A)   until   the   chamber   is

completely filled. Take care to fill all channels and

filling ports.

•

•

Leave at room temperature for at least 30 minutes.

Use the caps to decrease evaporation (step B).

Completely   aspirate   the   solution   using   a   cell

culture aspirator. Take care to aspirate from more

than  one  filling  port  to remove  as  much  liquid  as

possible (step C).

•

•

Wash with 80 μl ultra-pure water by injecting into

pipet ports 5 and 7 (step D). Please note that filling

the  chamber   with   water  or   washing  buffer   is

difficult as long as it is not completely dried. Don’t

mind air bubbles at this point.

Aspirate   exhaustively   by   using   a    cell   culture

aspirator. Take care to aspirate from more than one

filling  port  to  remove  as  much  liquid  as  possible.

(step E)

•

•

Repeat until the entire inside is washed.

Let    the   chamber    dry   at    room    temperature

overnight under sterile conditions.

Please   note   that   filling   the   chamber   with   cells   or

washing buffer is  difficult as long  as it is not  completely

dried.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 3 of 15

Application Note 14

b)  Observation Area Coating Protocol with Collagen IV

(equal to ibidi Cat. No. 80302)

This  protocol  is  used  to  coat  the  observation  area  only.

This  will  create  a  coating   barrier  so  cells  may  not  be

allowed to enter the large reservoirs. Only the observation

area and the side channels are coated with Collagen IV.

(μ-Slide   Chemotaxis   2D,   ibiTreat,   0.8  μg/cm  ,   6  μl,

2

60 μg/ml, 0.45 cm

2

)

•

•

•

•

Dilute the Collagen IV (e.g. BD Cat.-Nr. 356233) to

a concentration of 60 μg/ml using 0.05 M HCl.

Close  filling  ports  6  and  7  by  plugs  (not  shown).

Handle plugs with appropriate blunt tweezers.

Use a 20 μl pipet (e.g. Gilson P-20) and apply 6 μl

coating solution onto filling port 5 (step A).

Use the same pipet settings and aspirate air from

the opposite filling port 4. (step B) Press the pipet

tip directly into filling port 4. The coating solution

on  port  5  will  be  flushed  inside  filling  the  entire

channel  homogeneously.  Gently  aspirate  until  the

solution reaches the pipet tip.

•

•

•

•

Remove  the  plugs  and  cover  with  lids.  Leave  at

room temperature for 60 minutes (not shown).

Gently  close  filling  ports  6  and  7  by  plugs  (not

shown).

Put  10  μl  ultra-pure  water  on   top  filling  port 5

(step C). Don’t trap air bubbles.

Aspirate the same amount of liquid from cell inlet

adapter  5  as  shown   (step  C).  Repeat  this  step

twice.

•

•

Aspirate intensively all liquid from filling port 4 and

5 (step D).

Let    the   chamber    dry   at    room    temperature

overnight under sterile conditions.

After   coating,  the   chamber   needs  to   be   completely  dry.

Otherwise, the cell seeding protocol does not work correctly.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 4 of 15

Application Note 14

5.  Experiment

The   day   before   seeding   cells   and

conducting

the

experiment,

it

is

necessary to  place  all cell  media, the

μ-Slide,  and plugs  into  the incubator

for gas equilibration. Media  is put in a

slightly opened  vial. This  will prevent

the  medium inside  the  slide  and the

slide itself  from emerging  air bubbles

over the incubation time.

a)  Seeding cells

•

•

Unpack  the  μ-Slide  Chemotaxis  and   put  into  a

10 cm Petri dish with sterile, wet tissue around the

slide decreasing evaporation.

Prepare  your  cell  suspension  as  usual.  Use  cell

concentration  of  ca.   3  x  10   cells/ml.   High  cell

6

concentrations are needed due to the small height

of the observation area.

•

•

Close  filling  ports  6  and  7  by  plugs  (not  shown).

Handle plugs with appropriate blunt tweezers.

Use a 20 μl pipet (e.g. Gilson P-20) and apply 6 μl

cell suspension on top of filling port 5 (step A). Do

not inject directly.

•

Use the same pipet settings and aspirate air from

the  opposite   filling  port   4.  Press  the   pipet  tip

directly  into  filling  port  4.  The  cell  suspension  on

top of port 5 will be flushed inside filling the entire

channel  homogeneously.  Gently  aspirate  until  the

cell    suspension   reaches    the    pipet   tip.    The

hydrophilic  ibiTreat  surface  (TC-treated)   supports

this step so different surface coatings can behave

differently.

•

Leave  both  filling  ports  (4  and  5)  filled  with  cell

suspension.  Level  out  the  liquid  heights  in  both

filling ports like this:

Do not overfill or underfill both filling ports to avoid

leakage.  To  achive  this,  the  6 μl  seeding  volume

may have to be adjusted.

•

•

•

•

Directly after cell seeding, gently remove the plugs

from filling port 6 and 7 by tweezers (not shown).

Cover all filling ports with cultivation caps (step C).

Do not use the plugs during cell attachment.

Control  cell  density  and  cell  distribution  by  phase

contrast microscopy.

Incubate   the  slide   inside   a   sterile  and   humid

atmosphere  to   minimize   evaporation  until   cells

have attached. This should occur within 1-5 hours

depending  on  your  cell  type  and  surface.  Make

sure  evaporation  is  low  by  using  a  sterile  10  cm

Petri dish with extra wet tissue around the slide.

Observation area with recom-

mended cell density right after

seeding.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 5 of 15

Application Note 14

b)  Cell Attachment

•

Control cell attachment and morphology by phase

contrast    microscopy.    Control    evaporation    by

checking  the  amount  of  medium  in  filling  port  4

and 5.

After cell attachment (1 – 5 h) remove non-adherent cells

and old medium:

•

•

•

•

•

Remove  all  cultivation  caps  from  the  filling  ports

(not shown).

Gently  close  filling  ports  6  and  7  by  plugs  (not

shown).

Fill 10 μl cell free medium onto filling port 5. Don’t

trap air bubbles (step D).

Aspirate the same amount of liquid from cell inlet

adapter 4 as shown (step D).

Cells in observation area after

attachment.

Repeat this step twice.

This  washing  protocol  can  also  be  used  to  pretreat  the

cells before the chemotaxis experiment is started.

c)  Filling the Reservoirs

•

Transfer the plug from filling port 7 to filling port 4.

(not shown)

•

Fill one reservoir by gently injecting 45 μl cell-free

medium into filling port 5 (step E). In this step it is

crucial not to detach cells. This can be achieved by

very  gently   pipetting.  Slowly   inject  media   until

filling  port  7  is  completely  filled.  (This  is  a  very

critical step  because  cells might be flushed away

due to high flow rates.   The longer the filling takes

the   fewer    your    cells   are    stressed.   Approx.

30 seconds per reservoir are fine.)

•

•

Transfer the plug from filling port 6 to filling port 7.

(not shown)

Fill the other reservoir by gently injecting 45 μl cell-

free  medium  into  filling  port  5  (step  F).  (Keep  in

mind that all filling ports must be completely filled

but not overfilled, as shown.)

•

Now,  the  chamber  is  completely  filled  and  cells

grow  in  the  observation  area  only.  Control  your

cells under the phase contrast microscope.

d)  Apply a Chemoattractant in one Reservoir

Make   sure  the   two   open   filling  ports   5   and   6  are

completely filled like this:

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 6 of 15

Application Note 14

•

Apply 18 μl of chemoattractant onto filling port 6 as

shown.  Don’t  trap   air  bubbles.  Don’t  inject  the

chemoattractant directly (step G).

•

Aspirate  the  same  amount  of  liquid  (18 μl)  from

filling   port    5   (step    H).   This    will   flush    the

chemoattractant   into  the   large   reservoir.  Don’t

aspirate too fast as this might detach the cells.

Close  both  filling  ports  5  and  6  with  plugs.  Start

with  the   non-attractant  containing   filling  port   5

(step I). Don’t trap air bubbles.

After a short time the chemoattractant will diffuse

through the observation area and establish a linear

concentration   profile  over   the   cells.  See   § 10.

(fluorescence measurements) for details.

•

•

Please  visit  www.ibidi.com  for  movies  on  handling  and

applications withμ-Slide  Chemotaxis!

6.  Chemoattractant Concentrations and Recommended Experimental Setups

a)  Definition

For optimal results we recommend performing two control experiments by filling the chamber

completely with/without chemoattractant solution. That gives access to the question whether

a compound is influencing directed movement of cells (chemotaxis) and/or influencing random

migration (chemokinesis).

C100       Concentration   in    stable    gradient   equilibrium

(= max. concentration reaching the cells on one

edge of the observation area).

Capplied   Concentration  applied  into  the  chamber  with  a

volume of 18 μl. Capplied = C100 x 2.5

1)

Cbalanced  Balanced  chemoattracntant   concentration  after

infinite  time  calculated  with  the  total  chamber

volume of 80 μl. Cbalanced = Capplied x 18 μl / 80 μl.

C0

Solution    without   (or    low    concentration)   of

chemoattractant.

1)

In theory this value should be 40 μl / 18 μl = 2.22. Fluorescence measurements revealed

that  a   slightly  higher  concentration   is  needed   due  to  the   fact  that  in   equilibrium  the

chemotattractant containing reservoir already lost molecules for gradient establishment.

b)  Calculation Example

Chemoattractant concentration profile over cells between

100 μg/ml (=C100) and 0 μg/ml (=C0) in equilibrium.

C100

= 100 μg/ml

Capplied   = 100 μg/ml x 2.5 = 250 μg/ml

Cbalanced  = 250 μg/ml x 18 μl / 80 μl = 56.25 μg/ml

C0

= 0 μg/ml

The  used  solutions  are:  Capplied  for  +/-  experiment  and  C0

and Cbalanced  for the two control experiments (-/- and +/+).

Preparation of C100 is not necessary.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 7 of 15

Application Note 14

7.  Video Microscopy

Video microscopy is an absolutely necessary tool for ibidi’s μ-Slide Chemotaxis. Without, there

is no access and analysis of chemotaxis and migration effects.

•

Mount the μ-Slide Chemotaxis on the stage of your inverted microscope and observe

cell  movement  within  the  observation  areas.  Depending  on  the  cells´  velocity  one

frame  per  1-10  minutes  should  be  sufficient.  For  optimal   migration  data  of  slow

migrating cells we recommend starting the experiment immediately after filling in the

attractant and observe cell movement for 24 hours.  After initial experiments this time

can be varied for efficient use of experimental time.

•

•

•

•

Depending on your cells´ requirements heating and incubation devices are necessary.

Please contact us for more details on incubation systems right on the microscope.

After  mounting  the  slide  on  the  microscope  stage,   wait  20  min  for  temperature

equilibration.

Start  a  time-lapse  experiment  with  a  5x  or  10x  objective.  Higher  magnifications  are

possible but for optimal migration data it is necessary to track at least 20 – 40 cells.

After the time-lapse export your images as uncompressed single page .tif files.

Depending on your focus of interest, choose a magnification:

4x/5x

10x

- lowest magnification

- lowest resolution

- medium magnification

- medium resolution

- maximum overview

- maximum number of cells

- medium number of cells

40x/100x

- highest magnification

- maximum resolution

- smallest number of cells

20x

- medium magnification

- medium resolution

- medium number of cells

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 8 of 15

Application Note 14

8.  Tracking Cells

After  the  experiment,  it  is  necessary  to  track  the  cells with  the  appropriate  software.  We

recommend, for example, the ImageJ plugin, Manual Tracking. This plugin is able to quantify

the movement of objects between frames of a temporal stack.

•

•

ImageJ is available here: http://rsb.info.nih.gov/ij/

The Manual Tracking plug-in (including a PDF documentation) by Fabrice Cordelières,

Institut Curie, Orsay, France is available here:

http://rsbweb.nih.gov/ij/plugins/track/track.html

•

Download Manual_Tracking.class into the plugins folder of ImageJ on your computer

and restart ImageJ. Make sure you have administrator rights.

Quick guide for tracking with Manual Tracking:

1)   Import  a  movie  as  single  page  .tif  files,  by  using

„File/Import/Image_Sequence”.

2)   Open plugin “Manual Tracking”.

3)   Select “Add track”.

4)   Follow the first cell, through all of the time points,

by  clicking  on  the  cell’s  midpoint.  After  the  first

click, the software   creates the results table in an

extra window. This table is filled with the x/y data

of each cell, at each time point.

5)   Save the data table after tracking is completed.

6)   The  data  table  contains  all  tracked  cells  (=tracks)

and time points (=slices) with (x,y) positions.

•

•

For tracking, we recommend printing out the first image of each image stack before

the movie is watched for the first time. At least 40 cells are uniformly marked in order

to 1) ensure homogeneity and 2) avoid tracking the same cell twice. At least 30 cells

in the observation field need to be tracked during    the entire experiment. Some cells

will be lost through cell death, cell division, and cells leaving the observation field. We

do not recommend using such cells for data analysis.

You can also export “Overlay Dots & Lines” as a movie file (.avi).

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 9 of 15

Application Note 14

9.  Analyzing Chemotaxis

ibidi provides a free software tool for plotting and analyzing the tracked data. For a copy, go

to: http://www.ibidi.de/applications/ap_chemo.html

Here is a quick guide for data analysis:

1)   Import the data table from “Manual Tracking” (tab-

separated .xls file)

2)   Select the number of slices (=number of pictures

used  for  tracking).  The  number  of  slices  can  be

found  in  your  original  data  table  (“Show  original

data”).

3)   Calibrate the software by setting the x/y calibration

and the time interval. x/y calibration is the length of

one pixel in μm.

4)   Press “Apply settings”, after changing values and

parameters.

5)   Create trajectory plot, and then export as an image.

6)   Export values of center of mass, FMI, and Rayleigh test.

7)   Calculate  a  two-tailed  unpaired  Student  t-test  with  independent  variances  for  the

single values from 6). Compare the chemotaxis experiment +/- and the two control

experiments -/- and +/+ with each other to show statistical significance.

Example: Use the following command in Microsoft Excel.

=TTEST(A2:A4;B2:B4;2;3)

Here are some recommendations for presenting the results (in talks or publications):

•

•

•

•

•

Show the original movie (time lapse film of your cells).

Play the trajectory movie (time lapse film + overlaid cell trajectories).

Display the trajectory plot (graph).

Show the table or bar graph of center of mass, FMI, and Rayleigh test.

Include the results of the Student’s t-test.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 10 of 15

Application Note 14

10.Fluorescence Measurements

2

Fluorescence  measurements   with  Alexa488  (diffusion   coefficient  D=316 μm /s)  at   room

temperature (22 °C) revealed the concentration profile as linear over the 1 mm observation

area.  The  profile  is  established  after  1-8  h.  After  that  it  reaches  steady  state  (equilibrium).

Please note that there is no gradient inside the reservoirs. The profile is maintained linear for

over  48  h.  After  that  it  slowly  flattens  due  to  the  concentration  equilibration  between  the

reservoirs.

11.Different Chemoattractants and Time Dependencies

If you use a chemoattractant with different molecular weight    Mmolecule

(and diffusion constant

Dmolecule), different temperatures   T , and solutions with different dynamic viscosities    η  solution

,

diffusion  will  be  faster  or  slower.  The  following  example  (Alexa488  in  water  vs.  VEGF  in

endothelial  cell  growth  medium)  shows  how  to  compare  our  fluorescence  measurements

with different chemoattractants in different solutions.

Parameters:

ηAlexa488  (22°C)  = 1·10 -3

kg/s·m

ηEC growth medium (37°C)  = 0.75·10 -3

kg/s·m

T1 =37°C=310K

T2 =22°C=295K

To simplify matters the diffusing molecules are assumed to be globated. The hydrodynamic

radius is therefore dependent on the molecular weight only:

R h ≈  3  M  molecule

with:

MVEGF =45,000Da

MAlexa488=570Da

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 11 of 15

Application Note 14

Comparison:

T1

310K

k BT1

T1

ηECgrowthmedium(37°C)    3  MVEGF   =  0.75·10-3 kg/s·m3  45,000Da

DVEGF

DAlexa488

=  6πη1   R

η1 R  h1

T2

h1

=

=

k  T

T2

295K

B

2

6πη2  R h2

η2 R h2

ηAlexa488solution  (22°C)   3 MAlexa488

1·10-3 kg/s·m3  570 Da

DVEGF

= 0.33

DAlexa488

By  using  VEGF  in  endothelial  growth  medium  as  chemoattractant  at  37°C  the  diffusion  is

slower with factor 3. The concentration profile is build up slower with that factor. The diffusion

of different chemoattractants can be calculated identical.

12.Tips, Tricks and Troubleshooting

a)  Avoiding Air Bubbles

Air bubbles are 1) emerging over time due to not equilibrated media and slides or 2) injected

or  brought  into  the   chamber  by  the  user.  Both   types  will  disturb  the  diffusion  driven

concentration gradient by convection and must be avoided.

1)   Emerging  air  bubbles   over  time  can   be  avoided   by  equilibrating  all   media  and

equipment inside the incubator overnight. Media are put into a slightly opened vial.

Slides and plugs can be left in the sterile packaging.

2)   The second type of bubbles can be avoided by correct handling. Please avoid unfilled

filling  ports  and  trapped  air  bubbles  inside  the  pipet  tip!  Keep  in  mind  that  an  air

bubble might be trapped when closing an empty adapter by a plug. Always make sure

the filling ports are completely filled but not overfilled.

Avoid  unfilled  filling  ports and  air

bubbles inside the pipet tip!

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 12 of 15

Application Note 14

Accidentally empty filling ports can be refilled

by  a  10 μl  pipet  with  a  very  thin  pipet  tip.

Never  use  the   recommended  yellow  pipet

tips (see § 3) for this. An air bubble might be

trapped.

b)  Removing Air Bubbles

Air  bubbles  in  channels  and  reservoirs  can  often  be  removed  by  the  following  procedure.

Gently empty the reservoir containing the air bubble from the opposite side. Aspirate until the

atmospheric  air  reaches  the  air  bubble.  After  the  air  bubble  bursts  re-inject  the  aspirated

liquid. Always incline the slide so liquid is beneath the air phase.

Trapped air bubbles in filling ports can be removed by aspiration. Puncture them and aspirate

with a 10 μl pipet with a very thin pipet tip. To make sure there is enough liquid for a correctly

filled filling port, add a drop of liquid onto the trapped air bubble before the procedure.

c)  Inhomogeneous Cell Distribution

Inhomogeneous cell distribution, especially high cell densities

close  to  the  observation  area,  negatively  overlays  directed

migration  and  must  be  avoided.  Carefully  conduct  all  steps

following the protocol at § 5a to avoid spilling cell suspension

into the reservoirs. Never inject cell suspension directly into

filling ports.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 13 of 15

Application Note 14

d)  Chamber Fills not Properly

There are some general tips and tricks for filling the chamber:

¾   During filling by injection of liquid incline the slide in a way that liquid is beneath air.

Doing this air can escape trough the filling ports.

¾   Always use chambers that are completely dry.

Especially after coating the chambers need to be completely dry. Small liquid remains

might hinder proper handling.

¾   Use a pipet that is serviced routinely.

After some years of use pipets can loose pressure during pipetting. In normal work

they might perform great but there can be problems when used with μ-Slide

Chemotaxis.

e)  Cells do not Attach Properly

The chamber’s geometry is very special. Especially the cell seeding step is critical because

the volume is very low. Most problems with cell attachment can be avoided by checking the

following questions:

¾   Are the cells seeded in optimal growth medium?

Adherent cells will not attach properly when they are seeded in starvation medium.

Use normal growth medium with all necessary supplements like serum.

¾   Is the slide put into an extra humid chamber?

Evaporation is one of the most crucial issues during cell attachment. Make sure to

provide a maximum of humidity during cell attachment.

¾   Is the incubator frequently opened during cell attachment?

Frequent door openings must be avoided because this lowers humidity inside the

incubator drastically for a short time. A cell culture incubator that is not used by others

during cell attachment is perfect.

¾   Is the surface/coating suitable for the cell type you use?

Use a surface coating which is known to provide cell attachment for your cell type.

f)   Cells Are Flushed Away When Filling the Chamber

Most problems can be solved by checking the following points:

¾   Use adherent cells only.

¾   Fill the chamber with cell-free medium very carefully.

Depending on the filling speed there can be enormous shear forces during this step.

Control your cells under the phase contrast microscope after each pipetting step to

check during which step the problem occurs.

¾   Change the surface.

Since evaporation is critical during cell attachment it is favorable using a precoated

surface.  Coatings  decrease  attachment  time  because  cells  do  not  have  to  produce

their own extracellular matrix (ECM).

¾   Use μ-Slide Chemotaxis 3D

Our μ-Slide Chemotaxis 3D can also be used to conduct 2D experiments without the

use of gel. For this method, cells need to be slightly adherent to the surface of the

observation area only. More details can be found in our Application Note 17 in the

section 2D assay with slightly adherent cells.

g)  Scratches Inside the Observation Area

Scratches on the bottom of the slide can be due to putting the slide directly on a flat metal

surface  as  the  work  bench.  Avoid  draggling  the  slide  on  the  bench.  Always  use  the  slide

directly inside a 10 cm Petri dish.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 14 of 15

Application Note 14

Some scratches are part of the chamber itself. Unfortunately, those are normal in this slide

and unavoidable. These scratches are on the ceiling of the channel, not on the bottom, where

the cells are. Depending on the objective lens used (depth of focus) they are more or less

visible.

h)  Temperature Instabilities

Major changes in temperature during the experiment may cause convection thus disturbing

the diffusion driven gradient. Keep the temperature as stable as possible.

i)   Focus not Stable

Especially during time lapse experiments focus drift is an annoying effect. Focus stability is

mainly   influenced  by   mechanical   changes   and   temperature  instabilities.   Follow   these

recommendations to keep your cells in focus:

¾   Switch  on  all  components  (heating,  gas  incubation,  computer,  other  equipment)  at

least 60 min before starting the time-lapse recording.

¾   After you put the μ-Slide onto the microscope wait 20 min before starting a time lapse

experiment to achieve temperature (and immersion oil) equilibration.

¾   Keep the room temperature as stable as possible. Air conditioning should either be

working continuously or should be switched off.

¾   Do not change temperature during experiments. Avoid door/window openings as this

could change temperature rapidly.

¾   Eliminate   all  sources   of   mechanical  vibrations.   Use   a  damped   table   for   your

microscope.

¾   Use an autofocus system.

Please contact info@ibidi.com for more information.

Application Note 14

© ibidi GmbH, Version 2.6, Elias Horn, April 04

th

2012

Page 15 of 15