Instructions

µ–Dish 35mm,low

Theibidiproductfamilyiscomprisedofavarietyofµ–Slidesandµ–Dishes,whichhaveallbeen

designedforhigh–endmicroscopicanalysisofﬁxedorlivingcells.Thehighopticalqualityofthe

materialissimilartothatofglass,soyoucanperformallkindsofﬂuorescenceexperimentswith

uncompromisedresolutionandchoiceofwavelength.

Theµ–Dish35mm,lowallowsyoutoperformhighresolutionmicroscopyina35mmPetri–dishwith

7mmwalls. ThelowheightmakeshighnumericalaperturesofKo¨hlerilluminationpossibleand

provideslargeaccessformicromanipulation.   Thelidcanbeclosedtohinderevaporationduring

longtermexperiments.  ThelowheightmakeshighnumericalaperturesofKo¨hlerillumination

possibleandprovideslargeaccessformicromanipulation.Thelidcanbeclosedtohinder

Material

ShippingandStorage

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

ibidiµ–Slides,µ–Dishes,andµ–Platesaremadeofaplas-

tic that has the highest optical quality. The bottom ma-

terial exhibits extremely low birefringence and autoﬂuo-

rescence, similar to that of glass. Also, it is not possible

todetachthebottomfromtheupperpart.     Theµ–Slides,

µ–Dishes, and µ–Plates are not autoclavable, since they

areonlytemperature–stableupto80°C/175°F.Pleasenote

that gas exchange between the medium and incubator’s

atmosphere occurs partially through the polymer cover-

slip,whichshouldnotbecovered.

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

Conditions

Shippingconditions

Storageconditions

Ambient

RT(15-25°C)

ShelfLifeofDifferentSurfaces

ibiTreat,GlassBottom,ESS

36months

18months

4months

OpticalPropertiesibidiStandardBottom

Collagen,Poly-Lysine

Fibronectin

RefractiveindexnD(589nm)

Abbenumber

1.52

56

Thickness

No.1.5(180µm)

Material

microscopyplastic/

polymercoverslip

SurfaceandCoating

The µ–Dish 35mm,low is available with ibiTreat and un-

coated surface. The ibiTreat surface is a physical treat-

mentandoptimizedforadhesionofmostcelltypes.Many

celllinesaswellasprimarycellsweretestedforgoodcell

growth. Theuncoatedsurfaceisaveryhydrophobicsur-

faceandallowsnodirectcellgrowth.Itissuitableforspe-

ciﬁccoatingsorsuspensioncells.

Please note! The ibidi Standard Bottom is compatible

withcertaintypesofimmersionoilonly.    Alistofsuit-

ableoilscanbefoundonpage3.

If you like to establish a particular coating for your de-

mandswerecommendtotestyourcoatingprocedureon

uncoatedandibiTreatµ–Dishes,sincesomebiomolecules

adhere differently to hydrophobic or hydrophilic plastic

surfaces.

Geometry

Geometryoftheµ–Dish35mm,low

Diameterdish

35mm

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.  Prepareyour

µ–Dish,ibiTreatoruncoated.  Adjusttheconcentra-

tiontoacoatingareaof4.2cm2and400µl.

Volume

800µl

3.5cm2

Growtharea

Diametergrowtharea

Coatingareausing400µl

Heightwith/withoutlid

Bottommatchescoverslip

21mm

4.2cm2

• Apply400µlintothegrowtharea.    Makesurethat

theentirebottomiscoveredwithliquideasilytilting

orshakingtheµ–Dish.   Putonthelidandleaveat

roomtemperatureforatleast30minutes.

9mm/7mm

No.1.5

µ–Dish35mm,low

Page1

Version2015-05-12

Instructions

µ–Dish 35mm,low

• Aspiratethesolutionandwash.   Optionally,letdry

atroomtemperature.

• Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.

DetailedinformationaboutcoatingsisprovidedinAppli-

cationNote08

• Apply400µlcellsuspensionintotheinnerwellof

theµ–Dish.Avoidshakingasthiswillresultininho-

mogeneousdistributionofthecells.

Theµ–Dish 35mm,high isalsoavailablewithanelastically

supportedsurface(ESS)andaglassbottomsurface.Please

refertotheinstructionsfordetailedinformation.

• Aftercellattachmentaddadditionally400µlofpure

mediumtoensureoptimalgrowconditions.

ProtectionFilm

• Covertheµ–Dishwiththesuppliedlid.   Incubateat

37°Cand5%CO2asusual.

Removetheprotectionﬁlmbeforeusage!

Werecommendnottoﬁllmorethantheindicatedtotal

volume into the µ–Dish 35mm,low in order to avoid the

liquidcontactingthelid.

Thebottomoftheµ–Dishiscov-

ered with a ﬁlm to protect the

opticalqualityoftheplasticsur-

face. Pleasepullofftheprotec-

tionﬁlmbeforeusage!

Undemandingcellscanbeleftintheirseedingmediumfor

severaldaysandgrowtoconﬂuencethere.However,best

resultsareachievedwhenthemediumischangedevery

2–3days. Carefullyaspiratetheoldmediumandreplace

itbyupto800µlfreshmedium.

Theprotectionﬁlmofallµ–Dishesiscolor-coded:

Color

Surface

Variation

Tip:

Blue

ibiTreat

plain

Youcanstacktheµ–Dishestosavespaceinyourincu-

bator.Thiswillnotaffectcellgrowth.Werecommend

making batches with up to 6 µ–Dishes, due to sta-

bilityreasons. Placingtheµ–DishesintolargerPetri

dishessimpliﬁestransportandpreventsevaporation,

heat loss, and contamination when the incubator is

opened.

Red

ibiTreat

withGrid–500

withCulture-Insert

plain

Yellow

White

Black

Green

ibiTreat

uncoated

uncoated

uncoated

withGrid–500

withCulture–Insert

UsingTheLid

PreparationforCellMicroscopy

Toanalyzeyourcellsnospecialpreparationsarenecessary.

Cellscanbeobservedliveorﬁxeddirectlyintheµ–Dish

preferably on an inverted microscope. You can use any

ﬁxativeofyourchoice.Theµ–Dishmaterialiscompatible

withavarietyofchemicals,e.g.AcetoneorMethanol.Fur-

therspeciﬁcationscanbefoundatDueto

thethinbottomhighresolutionmicroscopyispossible.

1. openposition,easyopening

2. closeposition,forlongtermstudies,minimalevap-

oration

SeedingCells

For optimal results in ﬂuorescence microscopy and stor-

ageofstainedprobes,ibidiprovidesamountingmedium

optimized for µ–Dishes and µ–Slides (ibidi Mounting

Medium,50001).

Depending on your cell type, application of a 4–9 ×

104cells/mlsuspensionshouldresultinaconﬂuentlayer

within2–3days.

µ–Dish35mm,low

Page2

Version2015-05-12

Instructions

µ–Dish 35mm,low

µ–Dish 35mmSelectionGuide

MinimizingEvaporation

Usingtheµ–Dishwithaclosedlid,theevaporationinan

incubatorsystemwith37°Cand95%humidityisaround

1%perday.  Usingtheµ–Dishwithaclosedlidina37°C

heating system with low humidity (between 20% and

40%),theevaporationisaround10%perday.     Forreduc-

ingtheevaporationdownto1%perdayinallsystems,we

recommend sealing the lid with ibidi Anti–Evaporation

Oil(50051).

µ–Dish35mm,low

µ–Dish35mm,high

Low walls (7 mm) for  large  Highwalls(12mm)forallstan-

access to the cells.

Designed

dardapplications.

for micromanipulation and mi- Also available with glass bot-

croinjection.

tom,relocationgrid,andelastic

surface(ESS).

ImmersionOil

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

Company

Product

OrderingNumber

Zeiss

Zeiss

Leica

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

SelectedReferences

L.P.Frenzel,  Z.Abdullah, A.K.Kriegeskorte,  R.Dieterich, N.Lange, D.H.Busch,  M.Kronke, O.Utermohlen,  J.Hes-

cheler,andT.Saric.  RoleofNKG2D–ligandsandICAM–1inNKcell-mediatedLysisofMurineEmbryonicStemCellsand

EmbryonicStemCell–derivedCardiomyocytes.StemCells,2008.doi:10.1634/stemcells.2008-0528.

T.Quast,F.Eppler,V.Semmling,C.Schild,Y.Homsi,S.Levy,T.Lang,C.Kurts,andW.Kolanus.           CD81isEssentialfor

theFormationofMembraneProtrusionsandRegulatesRac1-ActivationinAdhesion-DependentImmuneCellMigration.

Blood,2011.doi:10.1182/blood-2010-12-326595.

K.Ro¨nneb¨aumer,U.Groß,andW.Bohne.   ThenascentparasitophorousvacuolemembraneofE.cuniculiisformedbyhost

celllipidsandcontainsporeswhichallownutrientuptake.EukaryoticCell,2008.doi:10.1128/ec.00004-08.

µ–Dish35mm,low

Page3

Version2015-05-12

Instructions

µ–Dish 35mm,low

µ–Dish 35mm,low

Family

µ–Dish35mm,low

Cat.No.

Description

Characteristics

80136

80131

µ–Dish 35mm,low ibiTreat: 35mm, lowwall(800µlvolume),    hydrophilic,sterilized

µ–Dish 35mm,lowUncoated:35mm,lowwall(800µlvolume),      hydrophobic,sterilized

#1.5polymercoverslip

#1.5polymercoverslip,tissueculturetreated

µ–Dish35mm,lowGrid-500

Cat.No.

Description

Characteristics

80156

µ–Dish 35mm,low ibiTreatGrid-500: 35mm,highwall(800µl     hydrophilic,sterilized

volume),#1.5polymercoverslip,tissueculturetreated,gridre-

peatdistance500µm

80151

µ–Dish35mm,lowUncoatedGrid-500:35mm,highwall(800µl      hydrophobic,sterilized

volume),#1.5polymercoverslip,gridrepeatdistance500µm

CultureInsertinµ–Dish35mm,low    ∗∗

Cat.No.

Description

Characteristics

80206

CultureInsertinµ–Dish35mm,low,ibiTreat:readytouse,tissue      hydrophilic,sterilized

culturetreated

80201

CultureInsertinµ–Dish35mm,low,uncoated:    readytouse,tis-   hydrophobic,sterilized

sueculturetreated

∗∗PleasealsorefertotheCulture-Insertinstructions.

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

µ–Dish35mm,low

Page4

Version2015-05-12