Instructions

µ–PlateAngiogenesis96well

Theibidiproductfamilyiscomprisedofavarietyofµ–Slidesandµ–Dishes,whichhaveallbeen

designedforhigh–endmicroscopicanalysisofﬁxedorlivingcells.     Thehighopticalqualityof

thematerialissimilartothatofglass,soyoucanperformallkindsofﬂuorescenceexperiments

withuncompromisedresolutionandchoiceofwavelength.

Theµ–PlateAngiogenesis96wellallowsyoutoperformtubeformationassaysongel,combined

withhighresolutionmicroscopyinastandardmulti–wellformat.

Material

SingleWellParameters

Volumeinnerwell

10µl

ibidiµ–Slides,µ–Dishes,andµ–Platesaremadeofaplas-

tic that has the highest optical quality. The bottom ma-

terial exhibits extremely low birefringence and autoﬂuo-

rescence, similar to that of glass. Also, it is not possible

todetachthebottomfromtheupperpart.     Theµ–Slides,

µ–Dishes, and µ–Plates are not autoclavable, since they

areonlytemperature–stableupto80°C/175°F.Pleasenote

that gas exchange between the medium and incubator’s

atmosphere occurs partially through the polymer cover-

slip,whichshouldnotbecovered.

Diameterinnerwell

Depthinnerwell

4mm

0.8mm

70µl

Volumeupperwell

Diameterupperwell

Growthareainnerwell

Coatingareausing10µl

5mm

0.125cm2

0.23cm2

AccuracyoftheBottom

Innerwellﬂatness

Wholeplateﬂatness

Bottom

±5µm

OpticalPropertiesibidiStandardBottom

±25µm

RefractiveindexnD(589nm)

Abbenumber

1.52

ibidiStandardBottom

56

Theµ–PlateAngiogenesis96wellmeetsallimportantval-

uesoftheANSI/SBSStandards(1-2004,2-2004,3-2004and

4-2004).

Thickness

No.1.5(180µm)

Material

microscopyplastic/

polymercoverslip

Please note! The ibidi standard bottom is compatible

withcertaintypesofimmersionoilonly.    Alistofsuit-

ableoilscanbefoundonpage3.

Geometry

The µ–Plate Angiogenesis 96well provides standard ge-

ometryandnumbering(A-H,1-12).  Thebottomoftheµ–

PlateAngiogenesis96wellprovidesahighaccuracy.

Dimensionsoftheµ–PlateAngiogenesis96wellinmm

Length

127.7

85.5

16.5

14.4

9.0

±0.2

±0.2

±0.4

±0.4

±0.1

Width

Heightwithlid

Heightwithoutlid

Welltowelldistance

µ–PlateAngiogenesis96well

Page1

Version2014-08-25

Instructions

µ–PlateAngiogenesis96well

ShippingandStorage

Non-gelbasedcoatingsarealsopossible.   Pleaseuse10µl

coatingsolutionandcalculatewithanareatobecoatedof

0.23cm2 perwell.  Furtherinformationaboutcoatingsis

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

providedin

SeedingCells

Conditions

Shippingconditions

Storageconditions

Ambient

In a tube formation assay cells are seeded on top of the

polymerizedgelmatrix:

RT(15-25°C)

ShelfLifeofDifferentSurfaces

ibiTreat,glassbottom,ESS

36months

18months

4months

Collagen,Poly-Lysine

Fibronectin

1. Trypsinizeandcountcellsasusual.    Dilutethecell

suspensiontothedesiredconcentration.Depending

onyourcelltype,werecommend1–3×105cells/ml.

Surface

2. Apply 70µl of the cell suspension into the upper

well.Donottouchthegelmatrixwiththepipettip.

Theµ–PlateAngiogenesis96wellisavailablewithibiTreat

surface. TheibiTreatsurfaceisaphysicaltreatmentand

optimizedforadhesionofmostcelltypes.Manycelllines

aswellasprimarycellsweretestedforgoodcellgrowth.

3. Cover the µ–Plate Angiogenesis 96well with the

suppliedlid.Incubateat37°Cand5%CO2asusual.

4. Conductyourexperiment.

Aspeciﬁccoatingoftheµ–PlateAngiogenesis96wellcan

bedoneyourselffollowingtheprocedureinsection”Coat-

ing”.

5. Depending on the cell type, medium exchange is

necessaryevery1–2days.  Carefullyaspiratetheold

mediumandreplaceitby70µlfreshmedium.

Foradetailedprotocolpleasereferto

and

RemovetheProtectionFilmbeforeUsage

Thebottomoftheµ–PlateAngiogenesis96welliscovered

withaﬁlmtoprotecttheopticalqualityoftheplasticsur-

face.Pleasepullofftheprotectionﬁlmbeforeusage!

Furtherinformationabouttheoptimizationofexperimen-

talparametersanddataanalysisisprovidedin

Coating

Tip:

Airbubblesinthegelcanbereducedbyequilibrating

theµ–PlateAngiogenesis96wellbeforeusageinside

theincubatorovernight.

Intubeformationassaystheµ–PlateAngiogenesis96well

iscoatedwitha0.8mmthicklayerofgelmatrix.

1. Prepare your gel matrix according to the manufac-

turer’sprotocolorreference.

2. Fill the inner well with 10 µl liquid gel. Avoid air

bubbles.

Tip:

Forlessevaporationthereservoirsattheedgescanbe

ﬁlled with sterile water or agarose. Add agarose to

water or buffer solution (e.g. 0.1 g to 10 ml water).

Melt agarose solution using a microwave or boiling

waterbathandallowthesolutiontocoolto∼50°C.

3. Letthegelpolymerizeunderappropriateconditions.

4. Useassoonaspossible.

5. Ifstorageisneededﬁllsterilewaterinthereservoirs

attheedgesoftheplate.

µ–PlateAngiogenesis96well

Page2

Version2014-08-25

Instructions

µ–PlateAngiogenesis96well

Tip:

Incasebentgelsurfacesarecreated,    increaseorde-

creasetheamountofgelused,untilyougetﬂatand

evengels.

• Filltheinnerwellwithalowvolumeofthegelma-

trix,e.g. 8µl.  Seedcellsontopofthepolymerized

gel. Ifnecessarygentlyshaketheslidetomakethe

cellsslideintothemiddleofthewell.

Tip:

Youcanstacktheµ–PlatesAngiogenesistosavespace

inyouincubator.  Thiswillnotaffectcellgrowth.   We

recommendmakingbatcheswithupto6plates,due

tostabilityreasons.

• Filltheinnerwellwithﬁbroblastssuspendedinside

agelmatrix.Seedcellsontopofthepolymerizedgel.

Overlaythecelllayerwithmediumandincubatefor

invasionofthecellsintothegelmatrix.

ExperimentalSetups

Alternatively, the µ–Plate Angiogenesis 96well can be

usedforthefollowingassays:

• Culturecellswithoutagelmatrixdirectlyinthemi-

norwells.

PreparationforCellMicroscopy

Whengelmatricesareusedtheopticalqualityandtheuse

ofhighmagniﬁcationobjectivelensesmightberestricted.

Withoutanygelcellscanbeobservedliveorﬁxeddirectly

inthewellsonaninvertedmicroscope.    Youcanuseany

ﬁxativeofyourchoice.  Theplasticmaterialiscompatible

withavarietyofchemicals,e.g.AcetoneorMethanol.Fur-

ther speciﬁcations can be found at  Due

to the thin bottom of only 180 µm, high resolution mi-

croscopyispossible.

• Filltheinnerwellwithcellssuspendedinsideagel

matrix. Aftergelpolymerization,add70µlcell–free

mediumtoﬁlltheupperwell.

ImmersionOil

When using oil immersion objectives, use only the im-

mersion oils speciﬁed in the table. The use of a non–

recommendedoilcouldleadtothedamageoftheplastic

materialandtheobjective.

• Fill the inner well with a gel matrix and culture

piecesoftissueorcellspheroidsonit.

Company

Product

OrderingNumber

ibidi

Zeiss

Zeiss

Leica

ImmersionOil

(ibidi)50101

Immersol518F

ImmersolW2010

Immersionliquid

(Zeiss)444960

(Zeiss)444969

(Leica)11513859

• SandwichCellCulture:Filltheinnerwellwithagel

matrix.Seedcellsontopofthepolymerizedgeland

imbedthecellswith70µlgelintheupperwell.

µ–PlateAngiogenesis96well

Page3

Version2014-08-25

Instructions

µ–PlateAngiogenesis96well

µ–PlateAngiogenesis96wellFamily

Theµ–PlateAngiogenesis96wellfamilyisavailablewithdifferentsurfacesandformats.Seetablebelowforchoosingyour

µ–Slideandµ–PlateAngiogenesis,respectively.

µ–SlideAngiogenesis

Cat.No.

Description

81506

81501

81531

µ–SlideAngiogenesisibiTreat:#1.5polymercoverslip,tissueculturetreated,sterilized

µ–SlideAngiogenesisuncoated:#1.5polymercoverslip,hydrophobic,sterilized

µ–SlideAngiogenesisMicrodissection:PEN-membrane∗ ,sterilized

∗ThePENfoildoesnotﬁttostandardcoverslipthickness.

µ–PlateAngiogenesis96well

Cat.No.

Description

89646

µ–PlateAngiogenesis96wellibiTreat:   #1.5polymercoverslip,  tissueculturetreated,

sterilized

References

P.S.Chen,Y.W.Shih,H.C.Huang,andH.W.Cheng.       Diosgenin,aSteroidalSaponin,InhibitsMigrationandInvasion

ofHumanProstateCancerPC-3CellsbyReducingMatrixMetalloproteinasesExpression.     PLoSONE,2011.  doi:10.1371/

journal.pone.0020164.

J.Searle,M.Mockel,S.Gwosc,S.A.Datwyler,F.Qadri,G.I.Albert,F.Holert,A.Isbruch,L.Klug,andD.N.Muller.Heparin

StronglyInducesSolublefms-LikeTyrosineKinase1ReleaseInVivoandInVitro.      Arteriosclerosis,Thrombosis,andVascular

Biology,2011.doi:10.1161ATVBAHA111.237784.

E. Vo, D. Hanjaya-Putra, Y. Zha, S. Kusuma, and S. Gerecht.  Smooth-Muscle-Like Cells Derived from Human Em-

bryonic Stem Cells Support and Augment Cord-Like Structures In Vitro.  Stem Cell Reviews and Reports , 2010.

doi:

10.1007/s12015-010-9144-3.

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

µ–PlateAngiogenesis96well

Page4

Version2014-08-25