Important Handling Information

for the µ-Slide Chemotaxis

2D

READ THIS BEFORE USE

Read this information sheet in its entirety BEFORE working with the µ-Slide

2D

Chemotaxis   . The following points are essential for successful handling.

1. Documents

•

•   Read Application Note 14

•   Watch the handling video:

2. Pipet Tips

Due to the filling ports’ design, this slide is only compatible with the following pipet tips. Using

other tips will skew your results.

•   Greiner Bio-One 739621, 739280, 739290,  772288

•   Axygen T-200-C, TR-222-C,  TR-222-Y

•   Starlab TipOne RPT S1161-1800

•   Sorenson BioScience MulTi Fit Tip 10590, 15320, 15330

3. Degassing Slides

Place slides and medium in an incubator before use to ensure that all gasses have been removed. Gas

bubbles will ruin your assay. You may leave the slide enclosed in the sterile package when doing this.

4. Humidity and Evaporation

Evaporation can  occur due to the  small volume of  the slides. It  is imperative to maintain  high

humidity during the attachment of your cells in order to avoid evaporation.

www.

.com

IMPORTANT HANDLING INFORMATION

Read this information sheet carefully BEFORE using the ibidi µ-Slide Chemotaxis

2D

General

Register for a 2 day

Laboratory Course at ibidi

Munich / Germany:

•   Read and directly follow instructions provided in Application Note 14.

•   Watch the handling video on www.ibidi.com.

http://ibidi.com/events/

practical-courses

•   The µ-Slide Chemotaxis

2D

is designed for adherent cells only.

•   The assay is not an endpoint measurement, thus cell migration

and chemotaxis must be followed by video microscopy. Go to

www.ibidi.com to download our Chemotaxis and Migration Tool.

•   Contact us at info@ibidi.com for support, questions, or comments.

Pipet Tips and Pipets

•   You must use the recommended 20-200 µl beveled pipet tips

when preparing the slide.

•   Make sure that your pipet is working properly and maintain

constant pressure during operation. Use a pipet which is serviced

and calibrated routinely.

•   Make sure to fill the pipet adapters on the slide fully; avoid air

bubbles inside the pipet tip.

Gas Equilibration

•   Incubate all media, slides and slide caps/plugs  overnight for gas

and temperature equilibration.

•   Loosen screw top on media container to ensure gas escape. Do

not remove completely in order to maintain sterility.

•   The µ-Slide Chemotaxis

2D

and slide plugs/caps  come in a sterile

package. This can be put directly into your incubator without

opening to maintain sterility.

Evaporation

•   To avoid evaporation during cell attachment, place the µ-Slide

2D

into a 10 cm Petri Dish and surround completely by

wet tissue. Use sterile tissue and water.

Chemotaxis

•   Do not open your incubator on a frequent basis during cell attachment.

Handling

•   Use the provided practice kit with colored water to practice pipetting.

•   Keep in mind that air has to escape when filling a chamber. Always

inject into a pipet adapter or channel which is already filled.

•   If coating the slides yourself, the chambers must be allowed to

dry completely.

ibidi GmbH

Tel.:

Fax:

E-Mail:  info@ibidi.de

+49 89 / 520 46 17 - 0

+49 89 / 520 46 17 - 59

For more information on our

Am Klopferspitz 19

82152 Planegg/Martinsried

Germany

products, and to download our

full catalog, please visit us at:

www.ibidi.com

© ibidi GmbH, V 2.0, 2015 / 08