Instruction Manual

Heated Lid & Heated  Plate

for 4μ-Slides

(Universal Fit)

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

2

©ibidi GmbH 2014, version 1.0, 2014-08-27

Table of contents

Safety Considerations ..............................................................................................................4

Specifications ...........................................................................................................................6

Compatibility.............................................................................................................................7

Preface .....................................................................................................................................8

Notices .....................................................................................................................................8

Limited Warranty......................................................................................................................9

Installation Requirements.......................................................................................................10

Installation and Connecting Cables ........................................................................................11

Maintenance and Operation...................................................................................................12

Waste Treatment ...................................................................................................................13

1

2

Working Principle ............................................................................................................14

Overview.........................................................................................................................15

2.1

Product Description ................................................................................................15

Components...........................................................................................................15

Heated Lid ..............................................................................................................16

Heated Plate with Insert.........................................................................................17

Sterilization and Cleaning........................................................................................18

Setting into Operation.............................................................................................18

2.2

2.3

2.4

2.5

2.6

3

4

Temperature Calibration..................................................................................................19

3.1

3.2

Calibration Routine .................................................................................................19

Influence of Surrounding Temperature and Airflow................................................21

Troubleshooting ..............................................................................................................22

4.1

4.2

4.2.1

Focus Not Stable ....................................................................................................22

Evaporation Too High..............................................................................................22

Silicone Oil Procedure (ibidi’s Anti-Evaporation Oil)............................................22

5

WEEE/RoHS Compliance Statement ..............................................................................23

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

3

Safety Considerations

To  ensure  operation  safety,   the  ibidi  heating  devices  must  be   operated  correctly  and

maintained  according  to  a  regular  schedule.  Carefully  read  to  fully  understand  all  safety

precautions before operating the instrument. Please take a moment to understand what the

signal words WARNING!, CAUTION and NOTE mean in this manual.

Safety symbols

WARNING!

CAUTION

A WARNING! indicates a potentially hazardous situation that, if

not avoided, could result in serious injury or even death.

A CAUTION indicates a potentially hazardous situation that, if

not avoided, may result in minor or moderate injury. It may also

be   used   alert   against    damaging   the   equipment   or   the

instrument.

Do not proceed beyond a WARNING! or CAUTION notice until

you  understand  the  hazardous  conditions  and  have  taken  the

appropriate steps.

NOTE

A  NOTE  provides  additional  information  to  help  the  operator

achieve optimal instrument and assay performance.

READ MANUAL label. This label indicates that you have to read

the manual before using the instrument.

This symbol indicates important notes.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

4

©ibidi GmbH 2014, version 1.0, 2014-08-27

Nomenclature

TEMPERATURE

CONTROLLER

HEATED LID

HEATED PLATE

WARNING! Only operate the ibidi heating devices with an ibidi Temperature Controller.

WARNING! Only operate the ibidi heating devices with the supplied cables and plugs. Using

other cables, etc. poses a risk of electrical shock and fire.

WARNING! Do not operate the ibidi heating devices with substances, or under conditions,

that can cause a risk of explosion, implosion or the release of gases. Only operate the ibidi

heating devices with aqueous solutions.

CAUTION Ensure that the power plug is well accessible. The ibidi heating devices must be

installed in a way that they do not hinder the access to the power plug.

CAUTION Some accessible parts of the plate and lid can reach temperatures up to 55°C.

Avoid  touching  the  temperature-controlled  parts  of  the  system  when  you  have  set  the

temperature controller to high temperatures.

CAUTION If the glass plate of the heated lid breaks due to a mechanical impact, the shards

can lead to injuries if handled.

CAUTION Be aware that when the system is switched on, 10 V AC voltage is applied to the

underside of the glass plate. Do not touch the underside or put it in contact with anything

conductive.  This  could  cause  a  short  circuit  that  may  destroy  the  temperature  controller

and/or the lid.

CAUTION Only use the heating devices in dry rooms, and do not use the heating devices in

the cold room.

CAUTION Use the devices only for cell experiments with aqueous solutions.

CAUTION    Do    not     use    the    heating    devices     with    hazardous    substances    or

substances/materials that pose a risk of infections.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

5

Specifications

Supply Voltage

Heated Lid

Heated Plate

10 V / 50 Hz  - 10 V / 60 Hz

12 V / 50 Hz – 12 V / 60 Hz

Surrounding Conditions

Storage Temperature

Operating Temperature

Humidity

-5 – 50°C

15 - 40°C (Indoor use only)

80 % relative humidity up to 31°C,

30 % relative humidity up to 40°C

Operating Altitude

max. 2000 m

Temperature Control Range

Ambient temperature to +55°C

Net Dimensions (hxwxd)

Heated Plate with Lid

25.5 x 125.5 x 85.5 mm³

210 g

Weight

Heated Plate with Lid

Recommended  Temperatures

before Calibration

(see 3.1 Calibration Routine)

Heated Lid

Heated Plate  (Channel 2)

(Channel 1)

42°C

38°C

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

6

©ibidi GmbH 2014, version 1.0, 2014-08-27

Compatibility

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

7

Preface

This  manual  is  your  guide  to  using  the  ibidi  Heated  Lid  &  Heated  Plate  for  cell  culture

experiments  on   an  optical   microscope.   It  instructs   first-time  users   how  to   use  the

instrument, and serves as a reference for experienced users.

Before  using  the  ibidi  heating  devices,  please  read  this  instruction  manual  carefully,  and

make sure that the contents are fully understood. This manual should be easily accessible to

the operator at all times during instrument operation. When not using the devices, keep this

manual in a safe place. If this manual gets lost, order a replacement from ibidi GmbH.

Notices

1.   ibidi shall not be held liable, either directly or indirectly, for any consequential damage

incurred as a result of product use.

2.   Prohibitions on the use of ibidi software

•    Copying software other than for backup

•    Transfer or licensing of the right to use software to a third party

•    Disclosure of confidential information regarding software

•    Modification of software

•    Use  of  software  on  multiple  workstations,  network  terminals,  or  by  other

methods

3.   The  contents  of  this  manual   are  subject  to  change  without   notice  for  product

improvement.

4.   This manual is considered complete and accurate at the time of publication.

5.   This manual does not guarantee the validity of any patent rights or other rights.

6.   If an ibidi software program has failed causing an error or improper operation, it may

be due to a conflict from another program operating on the notebook (PC). In this

case, take corrective action by uninstalling the conflicting product(s).

7.   ibidi is a registered trademark of ibidi GmbH in Germany and other countries.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

8

©ibidi GmbH 2014, version 1.0, 2014-08-27

Limited Warranty

Products sold by ibidi, unless otherwise specified, are under warranty for a period of one

year from the date of shipment, and should be delivered free of defects in materials and

workmanship. If any defects in the product are found during this warranty period, ibidi will

repair or replace the defective part(s) or product free of charge.

THIS WARRANTY DOES NOT APPLY TO DEFECTS RESULTING FROM THE FOLLOWING:

1

2

IMPROPER OR INADEQUATE INSTALLATION.

IMPROPER  OR  INADEQUATE  OPERATION,  MAINTENANCE,  ADJUSTMENT,  OR

CALIBRATION.

3

4

5

UNAUTHORIZED MODIFICATION OR MISUSE.

USE OF UNAUTHORIZED CELL CULTURE WARE.

USE  OF  CONSUMABLES,  DISPOSABLES,  AND  PARTS  NOT  SUPPLIED   BY  AN

AUTHORIZED IBIDI DISTRIBUTOR.

6

7

CORROSION DUE TO THE USE OF IMPROPER SOLVENTS, SAMPLES, OR DUE TO

SURROUNDING GASES.

ACCIDENTS BEYOND IBIDI’S CONTROL, INCLUDING NATURAL DISASTERS.

This warranty does not cover consumables like cell culture dishes, reagents, and the like.

The warranty for all parts supplied and repairs provided under this warranty expires on the

warranty  expiration  date  of  the  original  product.  For  inquiries  concerning  repair  service,

contact ibidi after confirming the model name and serial number of your ibidi system.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

9

Installation Requirements

To ensure operation safety, observe the following conditions:

1.   Only operate the ibidi heating devices with an ibidi Temperature Controller and with

the supplied power cable.

2.   Ensure that the power plug of the power supply is well accessible. The ibidi heating

devices must be installed in a way that they do not hinder the access to the external

power supply and its power plug.

3.   Operate the heating devices in a temperature range of 15 – 40°C.

4.   Operate the heating devices in a humidity range of 0 – 80% RH up to 31°C and 30%

RH up to 40 °C.

5.   Operate the heating devices in an atmospheric pressure range of 800 – 1060 hPa.

6.   Do not operate the heating devices under conditions that pose a risk of explosion,

implosion, or the risk of any gas release.

7.   Avoid  strong  magnetic  fields  and  sources  of  high  frequency.  The  heating  devices

may  not  function  properly  when  operating  near  a  strong  magnetic  field  or  high

frequency source.

8.   Avoid  vibrations   from  vacuum   pumps,  centrifuges,   electric  motors,  processing

equipment and machine tools.

9.   Avoid dust and corrosive gas. Do not install the heating devices where they may be

exposed to dust, especially in locations exposed to outside air or ventilation outlets.

10. Do not install the heating devices in a location where they may be exposed to direct

sunlight.

11. Install the heating devices in a horizontal and stable position. This includes a table,

bench, or desk upon which the devices are installed.

12. Ensure that no air conditioner blows air directly onto the heating devices. This may

prevent stable conditions.

13. Install the heating devices in a location that allows easy access for maintenance.

NOTE:  The  above  conditions   do  not  guarantee  optimal  performance   of  these  heating

devices.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

10

©ibidi GmbH 2014, version 1.0, 2014-08-27

Installation and Connecting Cables

Connecting the ibidi heating devices to the ibidi Temperature Controller

Confirm that the power switch of the temperature controller is off (power switch is at the

right, rear side of the temperature controller).

Connect  the   heated  lid   and  the  heated   plate  to   the  connectors   at  the   rear  of  the

Temperature Controller. Connect each device to the connector labeled with the number of

the  channel  you  want  to  use.  (e.g.  Channel  “1”  for  the  heated  lid,  Channel  “2”  for  the

heated plate).

WARNING!  Only  operate  the  ibidi  heating  system  with  the  supplied  cables  and  plugs,

because not doing so may put you at risk of electrical shock and fire.

CAUTION Ensure that the power plug of the external power supply is well accessible. The

heating system must be installed in a way that it does not hinder the access to its power

plug.

Connect  the  temperature  controller  to  a  Notebook  by  using  the  supplied  network  cable.

(Only necessary when using a PC to control the Heating System.)

Switch on the temperature controller and the Notebook.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

11

Maintenance and Operation

Pay close attention to the heating devices operating environment and always keep it clean

so that the heating devices can be used in a stabilized condition over a long period of time.

Operating the ibidi Heating Devices

Use the devices only in the intended application described in this manual.

CAUTION Unintended use which is not in compliance with this manual, can lead to a failure

of the devices’ protective systems. This failure may cause personal injury or damage.

Cleaning the ibidi Heating Devices

Switch  off   the  temperature   controller   and  disconnect   the  heating   devices  from   the

controller.  To  clean  the  devices,  use  only  a  dry  cloth  or  a  cloth  dampened  with  water

(ultrapure) or common lab disinfection    solutions that are based on quaternary    ammonium

compounds (e.g., Barrycidal 36 or Pharmacidal).

CAUTION The use of ethanol or other types of organic solvents to clean the devices may

remove the devices’ paint.

Transporting the ibidi Heating Devices

First, switch off the temperature controller and then disconnect the heating devices from the

controller. Carry the devices carefully and avoid mechanical shocks.

Repairing the ibidi Heating Devices

Do not try to repair the devices by yourself. Contact the ibidi service to repair any damaged

components.

CAUTION The manual disassembling of the devices is not allowed. Manual disassembling

poses a risk of personal injury or damage to the devices. Contact ibidi service personnel if

there is need to disassemble the devices.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

12

©ibidi GmbH 2014, version 1.0, 2014-08-27

Waste Treatment

Waste  treatment  is  your   own  responsibility.  You  must  dispose   of  the  waste  using  a

company that specializes in waste recovery. Do not dispose of the ibidi heating devices in a

trash bin or at a public waste disposable site. For detailed information, please contact the

ibidi service.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

13

1   Working Principle

The heated plate and the heated lid are designed to keep cells at 37°C in microscopy slides.

The heated lid prevents condensation effects inside the entire incubation system. Humidity,

CO2-enriched,  and/or  O2-reduced/enriched  air  can  be  flushed  in,  which  makes  the  heated

chamber a fully controlled incubator — right on the microscope. This system enables the

user to perform live cell imaging.

The glass top of the lid and the bottom plate are actively heat-controlled. The heated plate

has the outer dimensions of a 96-well plate (127.5 mm x 85.5 mm).

Figure 1: Schematic Drawing of the ibidi Heating Devices

Figure 2: Heated Lid and Heated Plate (with Insert)

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

14

©ibidi GmbH 2014, version 1.0, 2014-08-27

2   Overview

2.1   Product Description

A brief overview is provided in this section, including the nomenclature for later descriptions

in this document.

Heating

Devices

Temperature

Controller

Optional: PC

Figure 3: Schematic Setup

2.2   Components

ibidi Heated Glass Lid, Universal Fit

ibidi Heated Plate in Multi-Well Format,

for 4 μ-Slide with Heating Insert

Calibration Set

(1 temperature sensor + 1 perforated μ-Slide

8well)

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

15

2.3   Heated Lid

The heated lid (127.5 mm x 85.5 mm) provides excellent optical quality, and also allows for

the  use  of  all  standard  microscopy  techniques,  including  differential  interference  contrast

(DIC).  The  upper  glass  part  is  electrically  heat  controlled.  The  heated  lid  fits  exactly  and

securely onto the heated plate.

Figure 4: Heated Lid

For gas incubation, the lid is equipped with inlets for (a) the gas flow and (b) the humidity

sensor.

Due to the height of the lid, we recommend using condensers with a working distance of

≥ 26 mm.

Heated plate, heated lid, and inserts are positioned by strong Neodym

Magnets! Please contact us for a non-magnetic system, in the case

that permanent magnetic fields will disturb your experiment!

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

16

©ibidi GmbH 2014, version 1.0, 2014-08-27

2.4   Heated Plate with Insert

The heated plate provides the base for the heating system. Its format (127.5 mm x 85.5

mm)  fits   every  multi-well  stage   or  multi-well  holder   (e.g.,  for  96-well  plates)   on  the

microscope. It is actively heat-controlled. The heated plate is the counterpart to the heated

lid.

The insert has two functions that both use the force of the integrated magnets. The first

function  is  to  hold  the μ-Slides  firmly  into  position,  so  as  to  avoid  displacement  during

microscope stage movements. The second is to create a tight contact of the  μ-Slides to the

metallic insert, which then maximizes the heat transfer and ensures the stable heat control

of the sample.

Figure 5: Heated Plate for 4μ-Slides  with Insert

While using the heated plate make sure that there are  μ-Slides placed

in all four slots. Even if only one slot is in use, the openings on the

bottom  of  the  plate  must  be  closed  to  guarantee  stable  incubation

conditions.

Heated plate, heated lid, and inserts are positioned by strong Neodym

Magnets! Please contact us for a non-magnetic system, in the case

that permanent magnetic fields will disturb your experiment!

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

17

2.5   Sterilization and Cleaning

The  Heated  Plate  and  the  Heated  Lid  do  not  have  to  be  sterile,  because  the  incubation

chamber has no direct contact with the cells and the cell media. If sterilization is necessary

for  some  reason,  we  recommend  using  isopropanol  (70%)  or  common  lab  disinfection

solutions based on quaternary ammonium compounds (e.g. Barrycidal 36, Pharmacidal).

All  Heating  Devices  can  be  cleaned.  We  recommend  using  ultrapure  water  for  cleaning.

Stains on the Heated Lid can be removed using isopropanol (70%) or lens/eyeglass cleaning

wipes.

2.6   Setting into Operation

Here is the procedure to start the operation of the heating devices in combination with the

ibidi Temperature Controller.

For the first power up, we recommend disconnecting the heating devices (e.g., the heated

lid and the heated plate). By doing this, you ensure that you do not start heating with the

wrong temperature settings. It is now possible to set the temperatures for the individual

channels. Afterwards, you can reconnect the heating devices.

Since the controller is preconfigured, we recommend connecting the heated lid to Channel 1

and  the  heated  plate  to  Channel  2.  For  subsequent  power   ups,  you  will  not  need  to

disconnect the heating devices.

Please  let  the  temperature  of  the  system  equilibrate  for  a

minimum of 30 minutes before you start your experiments.

Be  aware  that  when  the  system  is  switched  on,  10V  AC

voltage is applied to the underside of the glass plate. Do not

touch the underside or contact it with anything conductive.

This   could  cause   a   short  circuit   that   may  destroy   the

controller and/or the lid.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

18

©ibidi GmbH 2014, version 1.0, 2014-08-27

3   Temperature Calibration

The temperature at the position of the cells must be calibrated from time to time, due to the

following conditions: room temperature, air conditioning, heat loss via microscope, chamber

type or objective lens, and humidification. A calibration is recommended before beginning

the first experiment. The heat calibration should be done for each chamber type, in order to

control temperatures with an absolute accuracy of less than 1 °C.

For calibration, the ibidi Temperature Controller has a plug, located on the rear panel, for an

external temperature sensor. Before shipment, all controllers are run through an in-house

calibration  to  ensure  their  accuracy.   This  calibration  only  refers  to  the  actively   heated

devices, such as the plate and the lid. In most cases, the cell temperature will lie somewhat

below the temperature of the heated plate. To compensate for this effect, the heated plate

can be calibrated to the cell temperature that is sensed by the external temperature sensor,

which can be placed at the position of the cells.

The displayed temperatures are valid for the heating

devices  only.  The  temperature  near  the  cells  might

be different (lower in most cases)!

3.1   Calibration Routine

The recommended calibration routine is fairly simple. Follow these eight steps for the initial

temperature calibration:

1)   Plug the external temperature (e.g., thermo-couple type K) sensor into the rear panel

of the temperature controller. Select the correct input for the display of the sensor

(TC-K, see Instruction Manual Temperature Controller, Section A.2).

2)   Fill the μ-Slide with medium, then place the sensor as close as possible to the future

position of the cells. Use the prepared calibration set, which consists of a μ-Slide with

a perforated lid, as shown in Figure 6.

3)   Set all parameters to create conditions identical to those of the actual experiments,

such as room temperature, air conditioning, airflow, illumination, microscope settings,

etc.

4)   Set the temperature of the heated plate to the desired cell temperature and set the

temperature of the lid several degrees higher (recommended values: plate 38 °C, lid

42 °C). You can do this directly with the control buttons on the controller, or with the

TempControl software.

5)   Let the system equilibrate for 30 minutes.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

19

6)   In most cases the temperature shown by the external sensor will not match the value

of  the   plate.  In   this  case  increase   or  decrease   the  temperature  of   the  plate

(recommended in 0.5 °C steps for small deviations). Let the system equilibrate for

another 5 minutes.

Repeat this step until you reach the desired temperature at the external sensor.

7)   Note the adjusted set values in your documentation for future experiments.

8)   Check the temperature of the system from time to time.

It  is  recommended  to  perform  this   routine  for  both

insert types, and once every year.

It  is  recommended   to  perform  this  routine  when   a

change of temperature conditions has occurred.

Figure 6: Sensor Placement for Temperature Calibration

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

20

©ibidi GmbH 2014, version 1.0, 2014-08-27

3.2   Influence of Surrounding Temperature and Airflow

The  surrounding,  ambient   temperature  affects  the   temperature  inside  the   microscopy

chambers.  Devices,  such  as  computers  and  camera  controllers,  can  significantly  heat  up

small rooms. In this case, we recommend equilibrating the room temperature to the typical

experimental conditions at least 2-3 hours before calibrating.

Keep in mind that the temperature sensor can also measure room temperatures.

The airflow can enhance the effect of temperature changes in the vicinity of the incubation

chamber. In a case where the airflow (e.g., air conditioning, etc.) cannot be stopped, we

recommend protecting the microscope as much as possible from any strong airflow. Please

contact us for support.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

21

4   Troubleshooting

4.1   Focus Not Stable

Focus drift is a disturbing effect, especially during time-lapse experiments. Focus stability is

mainly   influenced  by   mechanical   changes   and  temperature   variations.   Follow   these

recommendations to keep your cells in focus:

•

•

•

•

•

Switch   on  all   components   (e.g.,  heating,   gas   incubation,  computer,   or   other

equipment) at least 60 minutes before starting the experiment.

After  you  put  the  sample  onto  the  microscope,  wait  20  minutes  before  starting  a

time-lapse experiment to achieve temperature and immersion oil equilibration. (*)

Keep the room temperature as stable as possible. Air conditioning should either be

working continuously or switched off.

Do   not  change   the   temperature  during   the   experiments.  Avoid   door/window

openings, as this could rapidly change the temperature.

Eliminate  all   sources   of  mechanical   vibrations.  Use   a  damped   table  for   your

microscope.

(*) In case the experiment needs to be started immediately, either after placing the sample

on  the  microscope  or  after  closing  the  lid,  we  recommend  controlling  the  focus  for  20

minutes. In the  first minutes after starting the experiment, temperature equilibration might

influence the focus/z-position of the cells.

4.2   Evaporation Too High

Depending on the incubating conditions, small   volumes might evaporate quickly, especially

during long-term experiments. We suggest using silicone oil (e.g. Anti-Evaporation Oil, ibidi,

50051)  to  decrease  evaporation.  If  you  have  an   actively  controlled  humidifying  device,

increase the set value of relative humidity.

4.2.1

Silicone Oil Procedure (ibidi’s Anti-Evaporation Oil)

Covering  the   medium  with  sterile  silicone   oil  prevents  all   evaporation  effects  and   is

compatible with cell culture. Please don’t use mineral oil, as this would be harmful to the

ibidi μ-Slides.

Equilibrate  oil   and  medium  inside   the  incubator  overnight.  This   step  helps  avoid   the

formation  of  air  bubbles,  and  pre-warms  the  solutions  to  37°C.  Afterwards,  fill  your  slide

with cells and medium. Cover the medium’s surface with an appropriate amount of silicone

oil. Don’t drip the oil directly onto the surface, but let it run down the edges by pressing the

pipette tip directly on the upper side of the reservoir. For example, when using the ibidi   μ-

Slide VI

0.4

, fill each reservoir with 30 μl cell-free medium and 30 μl of silicone oil.

Please contact us at info@ibidi.com for further troubleshooting help.

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

22

©ibidi GmbH 2014, version 1.0, 2014-08-27

5   WEEE/RoHS Compliance Statement

EU Directives WEEE and RoHS

To Our Valued Customers:

ibidi GmbH is committed to being a good corporate citizen. As part of that commitment, we

strive  to  maintain  an  environmentally  conscious  manufacturing  operation.  The  European

Union (EU) has enacted two directives, the first on product recycling (Waste Electrical and

Electronic  Equipment,  WEEE)  and  the  second  on  limiting  the  use  of  certain  substances

(Restriction on the use of Hazardous Substances, RoHS).

Two Categories of products covered by the WEEE Directive are currently exempt from the

RoHS Directive – Category 8, medical devices (with the exception of implanted or infected

products) and Category 9, monitoring and control instruments.

All of our products fall into either Category 8 or 9, and are currently exempt from the RoHS

Directive. We will, however, continue to monitor the application of the RoHS Directive in

relation to our products and will comply with any changes, as they apply.

Recycling is offered for our products that fall within the scope of the WEEE Directive. These

specific products, available for sale after August 13, 2005, will be identified with a “wheelie

bin” symbol.

When you see this symbol, do not dispose of this product with Municipal Waste. Special

Collection/Disposal is required.

Please contact us for feedback or further information.

ibidi GmbH

Am Klopferspitz 19

82152 Planegg-Martinsried

Germany

www.ibidi.com

info@ibidi.de

Tel.: +49 89 520 4617 0

Fax: +49 89 520 4617 59

Heated Lid & Heated Plate for 4 μ-Slides (Universal Fit), Instruction Manual

©ibidi GmbH 2014, version 1.0, 2014-08-27

23