Instruction Manual

Heated Lid & Heated  Plate

for Multi-Well Plates

(Nikon Ti-S-E or Ti-S-ER)

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

©ibidi GmbH, Version 1.1, January 27th 2015

2

Table of contents

Safety Considerations ..............................................................................................................4

Specifications ...........................................................................................................................6

Compatibility.............................................................................................................................7

Preface .....................................................................................................................................8

Notices .....................................................................................................................................8

Limited Warranty......................................................................................................................9

Installation Requirements.......................................................................................................10

Installation and Connecting Cables ........................................................................................11

Maintenance and Operation...................................................................................................12

Waste Treatment ...................................................................................................................13

1

2

Working Principle ............................................................................................................14

Overview.........................................................................................................................15

2.1

System Overview ...................................................................................................15

Components...........................................................................................................16

Heated Lid ..............................................................................................................17

Heated Plate with glass bottom .............................................................................17

Sterilization and Cleaning........................................................................................18

Setting into Operation.............................................................................................18

2.2

2.3

2.4

2.5

2.6

3

4

Temperature Calibration..................................................................................................19

3.1

3.2

Calibration Routine .................................................................................................19

Influence of Surrounding Temperature and Airflow................................................20

Troubleshooting ..............................................................................................................21

4.1

4.2

4.2.1

Focus Not Stable ....................................................................................................21

Evaporation Too High..............................................................................................21

Silicone Oil Procedure (ibidi’s Anti-Evaporation Oil)............................................21

5

WEEE/RoHS Compliance Statement ..............................................................................22

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

3

Safety Considerations

To ensure operation safety, the ibidi heating devices must be operated correctly and

maintained according to a regular schedule. Carefully read to fully understand all safety

precautions in this manual before operating the instrument. Please take a moment to

understand what the signal words WARNING!, CAUTION and NOTE mean in this manual.

Safety symbols

WARNING!

CAUTION

A WARNING! indicates a potentially hazardous situation which,

if not avoided, could result in serious injury or even death.

A CAUTION indicates a potentially hazardous situation which, if

not avoided, may result in minor or moderate injury. It may also

be used alert against damaging the equipment or the

instrument.

Do not proceed beyond a WARNING! or CAUTION notice until

you understand the hazardous conditions and have taken the

appropriate steps.

NOTE

A NOTE provides additional information to help the operator

achieve optimal instrument and assay performance.

READ MANUAL label. This label indicates that you have to read

the manual before using the instrument.

This symbol indicates important notes.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

©ibidi GmbH, Version 1.1, January 27th 2015

4

Nomenclature

TEMPERATURE

CONTROLLER

HEATED LID

HEATED PLATE

WARNING! Only operate the ibidi heating devices with an ibidi Temperature Controller.

WARNING! Only operate the ibidi heating devices with the supplied cables and plugs. Using

other cables, etc. poses a risk of electrical shock and fire.

WARNING! Do not operate the ibidi heating devices with substances, or under conditions

that can cause a risk of explosion, implosion or release of gases. Only operate the heating

devices with aqueous solutions.

CAUTION Some accessible parts of the plate and lid can reach temperatures up to 55°C.

Avoid touching the temperature-controlled parts of the system when you have set the

temperature controller to high temperatures.

CAUTION  The  glass  bottom  of  the  heated  plate  may  break  due  to  a  mechanical  impact.

Avoid contact between the glass bottom and the objective lens. Changing objective lenses

automatically may lead to a collision between the objective lens and the glass bottom. Never

change to an objective lens with a working distance smaller than 1 mm.

CAUTION If the glass plate of the heated lid or plate breaks due to a mechanical impact, the

shards can lead to injuries if handled.

CAUTION Be aware that when the system is switched on, 10 V AC voltage is applied to the

underside of the glass plates. Do not touch the underside or put it in contact it with anything

conductive. This could cause a short circuit that may destroy the temperature controller

and/or the heating devices.

CAUTION Do only use the heating devices in dry rooms and do not use the heating devices

in the cold room.

CAUTION Turn off the power switch in the back of the chassis of the temperature

controller, when the heating devices are not in use.

CAUTION Use the devices only for cell experiments with aqueous solutions.

CAUTION Do not use the devices with hazardous substances or substances/materials that

pose a risk of infections.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

5

Specifications

Supply Voltage

Heated Plate

Heated Plate – Glass Bottom

Heated Lid

12 V / 50 Hz  - 12 V / 60 Hz

10 V / 50 Hz  - 10 V / 60 Hz

10 V / 50 Hz  - 10 V / 60 Hz

Surrounding Conditions

Storage Temperature

Operating Temperature

Humidity

-5 – 50°C

15 - 40°C (Indoor use only)

80 % relative humidity up to 31°C,

30 % relative humidity up to 40°C

Operating altitude

Max. 2000 m

Temperature Control Range

Temperature Deviation

(between wells of 96 w. plate)

Ambient temperature to +55°C

Max +/- 0.5°C

Net Dimensions (hxwxd)

Heated Plate

8 x 236 x 155 mm

3

Heated Lid

25 x 160 x 108 mm

3

Heated Plate with Lid

30 x 236 x 155 mm³

Weight

Heated Plate

Heated Lid

530 g

210 g

Recommended Temperatures

before Calibration

(see 3.1 Calibration Routine)

Heated Lid

Heated Plate

(Channel 1)    42°C

(Channel 2)    37°C

Heated Glass Bottom (Channel 3)    37°C

Restrictions for objective lenses   Working distance > 1 mm

Restrictions for condensers

Working distance > 30 mm

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

©ibidi GmbH, Version 1.1, January 27th 2015

6

Compatibility

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

7

Preface

This  manual  is  your  guide  to  using  the  ibidi  Heated  Lid  &  Heated  Plate  for  cell  culture

experiments  on   an  optical   microscope.   It  instructs   first-time  users   how  to   use  the

instrument, and serves as a reference for experienced users.

Before  using  the  ibidi  heating  devices,  please  read  this  instruction  manual  carefully,  and

make sure that the contents are fully understood. This manual should be easily accessible to

the operator at all times during instrument operation. When not using the devices, keep this

manual in a safe place. If this manual gets lost, order a replacement from ibidi GmbH.

Notices

1.   ibidi shall not be held liable, either directly or indirectly, for any consequential damage

incurred as a result of product use.

2.   Prohibitions on the use of ibidi software

•    Copying software other than for backup

•    Transfer or licensing of the right to use software to a third party

•    Disclosure of confidential information regarding software

•    Modification of software

•    Use  of  software  on  multiple  workstations,  network  terminals,  or  by  other

methods

3.   The  contents  of  this  manual   are  subject  to  change  without   notice  for  product

improvement.

4.   This manual is considered complete and accurate at the time of publication.

5.   This manual does not guarantee the validity of any patent rights or other rights.

6.   If an ibidi software program has failed causing an error or improper operation, it may

be due to a conflict from another program operating on the notebook (PC). In this

case, take corrective action by uninstalling the conflicting product(s).

7.   ibidi is a registered trademark of ibidi GmbH in Germany and other countries.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

8

©ibidi GmbH, Version 1.1, January 27th 2015

Limited Warranty

Products sold by ibidi, unless otherwise specified, are under warranty for a period of one

year from the date of shipment, and should be delivered free of defects in materials and

workmanship. If any defects in the product are found during this warranty period, ibidi will

repair or replace the defective part(s) or product free of charge.

THIS WARRANTY DOES NOT APPLY TO DEFECTS RESULTING FROM THE FOLLOWING:

1

2

IMPROPER OR INADEQUATE INSTALLATION.

IMPROPER  OR  INADEQUATE  OPERATION,  MAINTENANCE,  ADJUSTMENT,  OR

CALIBRATION.

3

4

5

UNAUTHORIZED MODIFICATION OR MISUSE.

USE OF UNAUTHORIZED CELL CULTURE WARE.

USE  OF  CONSUMABLES,  DISPOSABLES,  AND  PARTS  NOT  SUPPLIED   BY  AN

AUTHORIZED IBIDI DISTRIBUTOR.

6

7

CORROSION DUE TO THE USE OF IMPROPER SOLVENTS, SAMPLES, OR DUE TO

SURROUNDING GASES.

ACCIDENTS BEYOND IBIDI’S CONTROL, INCLUDING NATURAL DISASTERS.

This warranty does not cover consumables like cell culture dishes, reagents, and the like.

The warranty for all parts supplied and repairs provided under this warranty expires on the

warranty  expiration  date  of  the  original  product.  For  inquiries  concerning  repair  service,

contact ibidi after confirming the model name and serial number of your ibidi system.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

9

Installation Requirements

To ensure operation safety, observe the following conditions:

1.   Only operate the ibidi heating devices with an ibidi Temperature Controller and with

the supplied power cable.

2.   Ensure that the power plug of the power supply is well accessible. The ibidi heating

devices must be installed in a way that they do not hinder the access to the external

power supply and its power plug.

3.   Operate the heating devices in a temperature range of 15 – 40°C.

4.   Operate the heating devices in a humidity range of 0 – 80% RH up to 31°C and 30%

RH up to 40 °C.

5.   Operate the heating devices in an atmospheric pressure range of 800 – 1060 hPa.

6.   Do not operate the heating devices under conditions that pose a risk of explosion,

implosion, or the risk of any gas release.

7.   Avoid  strong  magnetic  fields  and  sources  of  high  frequency.  The  heating  devices

may  not  function  properly  when  operating  near  a  strong  magnetic  field  or  high

frequency source.

8.   Avoid  vibrations   from  vacuum   pumps,  centrifuges,   electric  motors,  processing

equipment and machine tools.

9.   Avoid dust and corrosive gas. Do not install the heating devices where they may be

exposed to dust, especially in locations exposed to outside air or ventilation outlets.

10. Do not install the heating devices in a location where they may be exposed to direct

sunlight.

11. Install the heating devices in a horizontal and stable position. This includes a table,

bench, or desk upon which the devices are installed.

12. Ensure that no air conditioner blows air directly onto the heating devices. This may

prevent stable conditions.

13. Install the heating devices in a location that allows easy access for maintenance.

NOTE:  The  above  conditions   do  not  guarantee  optimal  performance   of  these  heating

devices.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

10

©ibidi GmbH, Version 1.1, January 27th 2015

Installation and Connecting Cables

Connecting the ibidi heating devices to the ibidi Temperature Controller

Confirm that the power switch of the Temperature Controller is off (power switch is at the

right, rear side of the Temperature Controller).

Connect the heated lid and the heated plate to the connectors at the rear of the

Temperature Controller. Connect each device to the connector labeled with the number of

the channel you want to use. (e.g. channel “1” for the heated lid, “2” for the heated plate

and “3” for the glass bottom of the heated plate)

Mount the heating system into the Nikon Ti-S-E or Ti-S-ER motorized stage. Secure the

system by using the three screws delivered with the Nikon motorized stage.

To remove the heating system from the motorized stage, remove the three screws and use

the cable to pull out the system.

WARNING!  Only  operate  the  ibidi  heating  system  with  the  supplied  cables  and  plugs,

because not doing so may put you at risk of electrical shock and fire.

CAUTION Ensure that the power plug of the external power supply is well accessible. The

heating system must be installed in a way that it does not hinder the access to its power

plug.

Connect  the  temperature  controller  to  a  Notebook  by  using  the  supplied  network  cable.

(Only necessary when using a PC to control the Heating System.)

Switch on the temperature controller and the Notebook.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

11

Maintenance and Operation

Pay close attention to the heating devices operating environment and always keep it clean

so that the heating devices can be used in a stabilized condition over a long period of time.

Operating the ibidi Heating Devices

Use the devices only in the intended application described in this manual.

CAUTION Unintended use which is not in compliance with this manual, can lead to a failure

of the devices’ protective systems. This failure may cause personal injury or damage.

Cleaning the ibidi Heating Devices

Switch  off   the  Temperature  Controller   and  disconnect   the  heating  devices   from  the

controller.  To  clean  the  devices  use  only  a  dry  cloth  or  a  cloth  dampened  with  water

(ultrapure) or common lab disinfection    solutions that are based on quaternary    ammonium

compounds (e.g., Barrycidal 36 or Pharmacidal).

CAUTION The use of ethanol or other types of organic solvents to clean the devices may

remove the devices’ paint.

Transporting the ibidi Heating Devices

First, switch off the Temperature Controller and then disconnect the heating devices from

the controller. Carry the devices carefully and avoid mechanical shocks.

Repairing the ibidi Heating Devices

Do not try to repair the devices by yourself. Contact the ibidi service to repair any damaged

components.

CAUTION The manual disassembling of the devices is not allowed. Manual disassembling

poses a risk of personal injury or damage to the devices. Contact ibidi service personnel if

there is need to disassemble the devices.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

12

©ibidi GmbH, Version 1.1, January 27th 2015

Waste Treatment

Waste  treatment  is  your   own  responsibility.  You  must  dispose   of  the  waste  using  a

company that specializes in waste recovery. Do not dispose of the ibidi heating devices in a

trash bin or at a public waste disposable site. For detailed information, please contact the

ibidi service.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

13

1   Working Principle

The stage top incubator for multi-well plates consists of a heated plate with a heated glass

bottom, and also a heated lid that is designed to keep cells at 37°C in multi-well plates. The

heated lid prevents condensation effects inside the entire incubation system. Humidity, CO  2

enriched,  and/or  O2   reduced/enriched  air  can  be  flushed   in,  which  makes  the  heated

chamber  a  fully  controlled  incubator—right  on  the  microscope.  This  enables  the  user  to

perform live cell imaging.

The glass top of the lid, the glass bottom of the plate and the plate itself are actively heat-

controlled. The heated plate has the outer dimensions of the exchangeable centerpiece of

the Nikon Ti-S-ER motorized stage (236 mm x 155 mm).

Figure 1: Schematic Drawing of the ibidi Heating devices

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

14

©ibidi GmbH, Version 1.1, January 27th 2015

2   Overview

2.1   System Overview

Temperature

Controller

Heating

Devices

Optional:

Computer

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

15

2.2   Components

ibidi Heated Glass Lid for Multi-Well

Plates (Nikon Ti-S-E / Ti-S-ER)

ibidi Heated Plate with heated glass

bottom for Multi-Well Plates

(Nikon Ti-S-E / Ti-S-ER)

Calibration Set

(1 temperature sensor + 1 perforated μ-Plate

96well)

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

16

©ibidi GmbH, Version 1.1, January 27th 2015

2.3   Heated Lid

The heated lid (160 mm x 108 mm) provides excellent optical quality, and also allows for the

use of all standard microscopy techniques, including differential interference contrast (DIC).

The upper glass part is electrically heat controlled. The heated lid fits accurately onto the

heated plate.

For gas incubation, the lid is equipped with inlets for (a) the gas flow and (b) the humidity

sensor.

Due to the height of the lid, we recommend using condensers with a working distance of

> 30 mm.

2.4   Heated Plate with glass bottom

The heated plate provides the base for the heating devices. Its format (236 mm x 155 mm)

fits exactly the exchangeable centerpiece of the Nikon Ti-S-ER motorized stage. The glass

bottom and the plate itself are both actively heated. The heated plate is the counterpart to

the heated lid.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

17

2.5   Sterilization and Cleaning

The  Heated  Plate  and  the  Heated  Lid  do  not  have  to  be  sterile,  because  the  incubation

chamber has no direct contact with the cells and the cell media. If sterilization is necessary

for  some  reason,  we  recommend  using  isopropanol  (70%)  or  common  lab  disinfection

solutions based on quaternary ammonium compounds. (e.g. Barrycidal 36, Pharmacidal)

All  Heating  Devices  can  be  cleaned.  We  recommend  using  ultrapure  water  for  cleaning.

Stains on the Heated Lid or the glass bottom of the Heated Plate can be removed using

isopropanol (70%) or lens/eyeglass cleaning wipes.

Make sure not to apply to much pressure when cleaning the glass bottom of the plate to

avoid damaging the glass.

2.6   Setting into Operation

Here is the procedure to start the operation of the heating devices in combination with the

ibidi Temperature Controller.

For the first power up, we recommend disconnecting the heating devices. By doing this, you

ensure that you do not start heating with the wrong temperature settings. It is now possible

to  set  the  temperatures  for  the  individual  channels.  Afterwards,  you  can  reconnect  the

heating devices.

Since the controller is preconfigured, we recommend connecting the heated lid to Channel 1

the heated plate to Channel 2 and the glass bottom of the heated plate to Channel 3. For

subsequent power ups, you will not need to disconnect the heating devices.

Please  let  the  temperature  of  the  system  equilibrate  for  a

minimum of 30 minutes before you start your experiments.

Be  aware  that  when  the  system  is  switched  on,  10V  AC

voltage is applied to the underside of the glass plate. Do not

touch the underside or contact it with anything conductive.

This   could  cause   a   short  circuit   that   may  destroy   the

controller and/or the lid.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

18

©ibidi GmbH, Version 1.1, January 27th 2015

3   Temperature Calibration

The temperature at the position of the cells must be calibrated from time to time, due to the

following conditions: room temperature, air conditioning, heat loss via microscope, chamber

type or objective lens, and humidification. A calibration is recommended before beginning

the first experiment. The heat calibration should be done for each chamber type, in order to

control temperatures with an absolute accuracy of less than 1°C.

For calibration, the ibidi Temperature Controller has a plug, located on the rear panel, for an

external temperature sensor. Before shipment, all controllers are run through an in-house

calibration  to  ensure  their  accuracy.   This  calibration  only  refers  to  the  actively   heated

devices, such as the plate and the lid. In most cases, the cell temperature will lie somewhat

below the temperature of the heated plate. To compensate for this effect, the heated plate

can be calibrated to the cell temperature that is sensed by the external temperature sensor,

which can be placed at the position of the cells.

The displayed temperatures are valid for the heating

devices  only.  The  temperature  near  the  cells  might

be different (lower in most cases)!

3.1   Calibration Routine

The recommended calibration routine is fairly simple. Follow these eight steps for the initial

temperature calibration:

1)   Plug  the   external  temperature  sensor   into  the   rear  panel   of  the  Temperature

Controller.

2)   Fill the multi-well plate you want to use with medium and place the sensor as close

as possible to the future position of the cells. Use the prepared calibration set, which

consists of a μ-Plate 96 well with a perforated lid, as shown in Figure 2 on page 20.

3)   Set all parameters to create conditions identical to those of the actual experiments,

such as room temperature, air conditioning, airflow, illumination, microscope settings,

etc.

4)   Set the temperature of glass bottom and heated plate to the desired cell temperature

and  set  the  temperature  of  the  lid  several  degrees  higher  (recommended  values:

glass 37°, plate 37 °C, lid 42 °C). You can do this directly with the control buttons on

the controller, or with the TempControl software.

5)   Let the system equilibrate for 30 minutes.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

19

6)   In most cases the temperature shown by the external sensor will not match the value

of the plate. In that case increase or decrease the temperature of the glass bottom of

the  plate  (recommended  in   0.5  °C  steps  for  small   deviations).  Let  the  system

equilibrate for another 5 minutes.

Repeat this step until you reach the desired temperature at the external sensor.

7)   Note the adjusted set values in your documentation for future experiments.

8)   Check the temperature of the system from time to time.

It is recommended to perform this routine once every year.

It is recommended to perform this routine when a change

of temperature conditions has occurred.

Figure 2: Sensor Placement for Temperature Calibration

3.2   Influence of Surrounding Temperature and Airflow

The  surrounding,  ambient   temperature  affects  the   temperature  inside  the   microscopy

chambers.  Devices,  such  as  computers  and  camera  controllers,  can  significantly  heat  up

small rooms. In this case, we recommend equilibrating the room temperature to the typical

experimental conditions at least 2-3 hours before calibrating.

Keep in mind that the temperature sensor can also measure room temperatures.

The airflow can enhance the effect of temperature changes in the vicinity of the incubation

chamber. In a case where the airflow (e.g., air conditioning, etc.) cannot be stopped, we

recommend protecting the microscope as much as possible from any strong airflow. Please

contact us for support.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

20

©ibidi GmbH, Version 1.1, January 27th 2015

4   Troubleshooting

4.1   Focus Not Stable

Focus drift is a disturbing effect, especially during time-lapse experiments. Focus stability is

mainly   influenced  by   mechanical   changes   and  temperature   variations.   Follow   these

recommendations to keep your cells in focus:

•

•

•

•

•

Switch on all components (heating, gas incubation, computer, other equipment) at

least 60 minutes before starting the experiment.

After  you  put  the  sample  onto  the  microscope,  wait  20  minutes  before  starting  a

time-lapse experiment to achieve temperature and immersion oil equilibration. (*)

Keep the room temperature as stable as possible. Air conditioning should either be

working continuously or switched off.

Do   not  change   the   temperature  during   the   experiments.  Avoid   door/window

openings, as this could rapidly change the temperature.

Eliminate  all   sources   of  mechanical   vibrations.  Use   a  damped   table  for   your

microscope.

(*) In case the experiment needs to be started immediately, either after placing the sample

on  the  microscope,  or  after  closing  the  lid,  we  recommend  controlling  the  focus  for  20

minutes. In the  first minutes after starting the experiment, temperature equilibration might

influence the focus/z-position of the cells.

4.2   Evaporation Too High

Depending on the incubating conditions, small   volumes might evaporate quickly, especially

during long-term experiments. We suggest using silicone oil (e.g. Anti-Evaporation Oil, ibidi,

50051)  to  decrease  evaporation.  If  you  have  an   actively  controlled  humidifying  device,

increase the set value of relative humidity.

4.2.1

Silicone Oil Procedure (ibidi’s Anti-Evaporation Oil)

Covering  the   medium  with  sterile  silicone   oil  prevents  all   evaporation  effects  and   is

compatible with cell culture. Please don’t use mineral oil, as this would be harmful to the

ibidi μ-Slides.

Equilibrate  oil   and  medium  inside   the  incubator  overnight.  This   step  helps  avoid   the

formation of air bubbles, and pre-warms the solutions to 37°C. Afterwards, fill your plate

with cells and medium. Cover the medium’s surface with an appropriate amount of silicone

oil. Don’t drip the oil directly onto the surface, but let it run down the edges by pressing the

pipette tip directly on the upper side of the reservoir.

Please contact us at info@ibidi.com for further troubleshooting help.

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

© ibidi GmbH, Version 1.1, January 27th 2015

21

5   WEEE/RoHS Compliance Statement

EU Directives WEEE and RoHS

To Our Valued Customers:

ibidi GmbH is committed to being a good corporate citizen. As part of that commitment, we

strive  to  maintain  an  environmentally  conscious  manufacturing  operation.  The  European

Union (EU) has enacted two directives, the first on product recycling (Waste Electrical and

Electronic  Equipment,  WEEE)  and  the  second  on  limiting  the  use  of  certain  substances

(Restriction on the use of Hazardous Substances, RoHS).

Two Categories of products covered by the WEEE Directive are currently exempt from the

RoHS Directive – Category 8, medical devices (with the exception of implanted or infected

products) and Category 9, monitoring and control instruments.

All of our products fall into either Category 8 or 9, and are currently exempt from the RoHS

Directive. We will, however, continue to monitor the application of the RoHS Directive in

relation to our products and will comply with any changes, as they apply.

Recycling is offered for our products that fall within the scope of the WEEE Directive. These

specific products, available for sale after August 13, 2005, will be identified with a “wheelie

bin” symbol.

When you see this symbol, do not dispose of this product with Municipal Waste. Special

Collection/Disposal is required.

Please contact us for feedback or further information.

ibidi GmbH

Am Klopferspitz 19

82152 Planegg-Martinsried

Germany

www.ibidi.com

info@ibidi.de

Tel.: +49 89 520 4617 0

Fax: +49 89 520 4617 59

Heated Lid & Heated Plate for Multi-Well Plates (Nikon Ti-S-E or Ti-S-ER)

22

©ibidi GmbH, Version 1.1, January 27th 2015