Heated Lid & Heated Plate,

Multi-Well Plates

for K-Frame Stage (160 mm x 110 mm)

Instructions

Version 1.0

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–   160x110mm

©ibidi GmbH, Version 1.0, May 28th 2015

2

)

2

Table of contents

Safety Considerations ...........................................................................................................   4

Specifications........................................................................................................................    6

Compatibility..........................................................................................................................    7

Preface..................................................................................................................................    8

Notices..................................................................................................................................    8

Limited Warranty ...................................................................................................................   9

Installation Requirements....................................................................................................    10

Installation and Connecting Cables .....................................................................................   11

Maintenance and Operation ................................................................................................   12

Waste Treatment.................................................................................................................    13

1

2

Working Principle .........................................................................................................   14

Overview......................................................................................................................    15

2.1

System Overview..................................................................................................15

Components .........................................................................................................16

Heated Lid ............................................................................................................17

Heated Plate with glass bottom.............................................................................17

Sterilization and Cleaning .....................................................................................18

Setting into Operation ...........................................................................................18

2.2

2.3

2.4

2.5

2.6

3

4

Temperature Calibration...............................................................................................   19

3.1

3.2

Calibration Routine ...............................................................................................19

Influence of Surrounding Temperature and Airflow ...............................................20

Troubleshooting............................................................................................................   21

4.1

4.2

4.2.1

Focus Not Stable ..................................................................................................21

Evaporation Too High ...........................................................................................21

Silicone Oil Procedure (ibidi’s Anti-Evaporation Oil) .........................................  21

5

WEEE/RoHS Compliance Statement ...........................................................................  22

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

3

Safety Considerations

To ensure operation safety, the ibidi heating devices must be operated correctly and

maintained according to a regular schedule. Carefully read to fully understand all safety

precautions in this manual before operating the instrument. Please take a moment to

understand what the signal words WARNING!, CAUTION and NOTE mean in this manual.

Safety symbols

WARNING!

CAUTION

A WARNING! indicates a potentially hazardous situation which,

if not avoided, could result in serious injury or even death.

A CAUTION indicates a potentially hazardous situation which, if

not avoided, may result in minor or moderate injury. It may also

be used alert against damaging the equipment or the

instrument.

Do not proceed beyond a WARNING! or CAUTION notice until

you understand the hazardous conditions and have taken the

appropriate steps.

NOTE

A NOTE provides additional information to help the operator

achieve optimal instrument and assay performance.

READ MANUAL label. This label indicates that you have to read

the manual before using the instrument.

This symbol indicates important notes.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

4

Nomenclature

TEMPERATURE

CONTROLLER

HEATED LID

HEATED PLATE

WARNING! Only operate the ibidi heating devices with an ibidi Temperature Controller.

WARNING! Only operate the ibidi heating devices with the supplied cables and plugs. Using

other cables, etc. poses a risk of electrical shock and fire.

WARNING! Do not operate the ibidi heating devices with substances, or under conditions

that can cause a risk of explosion, implosion or release of gases. Only operate the heating

devices with aqueous solutions.

CAUTION Some accessible parts of the plate and lid can reach temperatures up to 55°C.

Avoid touching the temperature-controlled parts of the system when you have set the

temperature controller to high temperatures.

CAUTION  The glass  bottom  of the  heated plate  may  break due  to  a mechanical  impact.

Avoid contact  between the  glass bottom and  the objective  lens. Changing  objective lenses

automatically may lead to a collision  between the objective lens and the glass bottom. Never

change to an objective lens with a working distance smaller than 1 mm.

CAUTION If the glass plate of  the heated lid or plate breaks due to a mechanical impact,  the

shards can lead to injuries if handled.

CAUTION Be aware that when the system is switched on, 10 V AC voltage is applied to the

underside of the glass plates. Do not touch the underside or put it in contact it with anything

conductive. This could cause a short circuit that may destroy the temperature controller

and/or the heating devices.

CAUTION Do only use the heating devices in dry rooms and do not use the heating devices

in the cold room.

CAUTION Turn off the power switch in the back of the chassis of the temperature controller,

when the heating devices are not in use.

CAUTION Use the devices only for cell experiments with aqueous solutions.

CAUTION Do not use the devices with hazardous substances or substances/materials that

pose a risk of infections.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

5

Specifications

Supply Voltage

Heated Plate

Heated Plate – Glass Bottom

Heated Lid

12 V / 50 Hz  - 12 V / 60 Hz

10 V / 50 Hz  - 10 V / 60 Hz

10 V / 50 Hz  - 10 V / 60 Hz

Surrounding Conditions

Storage Temperature

-5 –  50°C

Operating Temperature

Humidity

15 - 40°C (Indoor use only)

80 % relative humidity up to 31°C,

30 % relative humidity up to 40°C

Operating altitude

Max. 2000 m

Temperature Control Range

Temperature Deviation

(between wells of 96 w. plate)

Ambient temperature to +55°C

Max +/- 0.5°C

Net Dimensions (hxwxd)

Heated Plate

9 x 160 x 110 mm

3

Heated Lid

25 x 160 x 108 mm

3

Heated Plate with Lid

30 x 236 x 155 mm³

Weight

Heated Plate

Heated Lid

220 g

210 g

Recommended Temperatures

before Calibration

(see 3.1 Calibration Routine)

Heated Lid

Heated Plate

(Channel 1)   42°C

(Channel 2)   37°C

Heated Glass Bottom (Channel 3)    37°C

Requirements

lenses

for

objective  > 1 mm working distance for coverslip bottom plates,

e.g. ibidi’s µ-Plate 96 Well

> 2 mm working distance for thick bottom plates

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

6

©ibidi GmbH, Version 1.0, May 28th 2015

Compatibility

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

7

Preface

This  manual  is your  guide  to  using  the  ibidi Heated  Lid  &  Heated  Plate  for  cell culture

experiments on an optical microscope.  It instructs first-time users how to use  the instrument,

and serves as a reference for experienced users.

Before  using the  ibidi  heating  devices, please  read  this  instruction manual  carefully,  and

make sure that the contents are fully understood.  This manual should be easily accessible to

the operator at  all times during instrument  operation. When not using  the devices, keep this

manual in a safe place. If this manual gets lost, order a replacement from ibidi GmbH.

Notices

1.   ibidi shall not be held liable, either directly or indirectly, for any consequential damage

incurred as a result of product use.

2.   Prohibitions on the use of ibidi software

    Copying software other than for backup

    Transfer or licensing of the right to use software to a third party

    Disclosure of confidential information regarding software

    Modification of software

    Use  of  software  on  multiple   workstations,  network  terminals,  or  by  other

methods

3.   The  contents  of   this  manual  are   subject  to  change   without  notice  for   product

improvement.

4.   This manual is considered complete and accurate at the time of publication.

5.   This manual does not guarantee the validity of any patent rights or other rights.

6.   If an ibidi software program  has failed causing an  error or improper operation, it  may

be due  to  a conflict  from another  program operating  on  the notebook  (PC). In  this

case, take corrective action by uninstalling the conflicting product(s).

7.   ibidi is a registered trademark of ibidi GmbH in Germany and other countries.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

8

Limited Warranty

Products  sold by  ibidi, unless  otherwise specified,  are under  warranty  for a  period of  one

year  from the  date of  shipment,  and should  be delivered  free  of defects  in  materials and

workmanship. If  any defects  in the  product are  found during  this warranty  period, ibidi  will

repair or replace the defective part(s) or product free of charge.

THIS   WARRANTY    DOES   NOT   APPLY    TO   DEFECTS   RESULTING    FROM   THE

FOLLOWING:

1

2

IMPROPER OR INADEQUATE INSTALLATION.

IMPROPER OR  INADEQUATE OPERATION,  MAINTENANCE, ADJUSTMENT,  OR

CALIBRATION.

3

4

5

UNAUTHORIZED MODIFICATION OR MISUSE.

USE OF UNAUTHORIZED CELL CULTURE WARE.

USE OF  CONSUMABLES,  DISPOSABLES, AND  PARTS  NOT SUPPLIED  BY AN

AUTHORIZED IBIDI DISTRIBUTOR.

6

7

CORROSION DUE TO THE  USE OF IMPROPER SOLVENTS, SAMPLES,  OR DUE

TO SURROUNDING GASES.

ACCIDENTS BEYOND IBIDI’S CONTROL, INCLUDING NATURAL DISASTERS.

This warranty does not cover consumables like cell culture dishes, reagents, and the like.

The warranty  for all parts  supplied and repairs  provided under  this warranty expires  on the

warranty  expiration  date  of  the  original  product.  For inquiries  concerning  repair  service,

contact ibidi after confirming the model name and serial number of your ibidi system.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

9

Installation Requirements

To ensure operation safety, observe the following conditions:

1.   Only operate the  ibidi heating  devices with an  ibidi Temperature Controller  and with

the supplied power cable.

2.   Ensure that the  power plug of  the power supply  is well accessible.  The ibidi heating

devices must be  installed in a way that  they do not hinder  the access to the external

power supply and its power plug.

3.   Operate the heating devices in a temperature range of 15 –  40°C.

4.   Operate the heating devices in  a humidity range of 0 –  80% RH up to 31°C and 30%

RH up to 40 °C.

5.   Operate the heating devices in an atmospheric pressure range of 800 – 1060 hPa.

6.   Do not  operate the  heating devices  under  conditions that  pose a  risk of  explosion,

implosion, or the risk of any gas release.

7.   Avoid strong magnetic fields and sources of high frequency. The heating devices may

not function  properly when operating  near a strong  magnetic field or  high frequency

source.

8.   Avoid  vibrations   from   vacuum  pumps,   centrifuges,   electric  motors,   processing

equipment and machine tools.

9.   Avoid dust and  corrosive gas. Do  not install the  heating devices where they  may be

exposed to dust, especially in locations exposed to outside air or ventilation outlets.

10. Do not install  the heating devices in  a location where they  may be exposed to  direct

sunlight.

11. Install the  heating devices in  a horizontal and  stable position.  This includes a  table,

bench, or desk upon which the devices are installed.

12. Ensure that  no air conditioner  blows air  directly onto  the heating devices.  This may

prevent stable conditions.

13. Install the heating devices in a location that allows easy access for maintenance.

NOTE:  The  above   conditions  do  not  guarantee  optimal   performance  of  these  heating

devices.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

10

Installation and Connecting Cables

Connecting the ibidi heating devices to the ibidi Temperature Controller

Confirm that  the power  switch of  the Temperature  Controller is  off (power  switch is  at the

right, rear side of the Temperature Controller).

Connect the heated lid and the heated plate to the connectors at the rear of the Temperature

Controller. Connect each device to the connector labeled with the number of the channel you

want to use. (e.g. channel “1” for the heated lid, “2” for the heated plate and “3” for the glass

bottom of the heated plate)

Mount the heating system into the motorized stage of your inverted microscope.

To remove the heating system from the motorized stage, use the cable to pull out the

system.

WARNING!  Only  operate  the  ibidi  heating  system  with  the  supplied  cables  and  plugs,

because not doing so may put you at risk of electrical shock and fire.

CAUTION Ensure  that the power  plug of  the external power  supply is well  accessible. The

heating system  must be  installed  in a  way that  it does  not hinder  the access  to its  power

plug.

Connect the temperature controller to a Notebook by using the  supplied network cable. (Only

necessary when using a PC to control the Heating System.)

Switch on the temperature controller and the Notebook.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

11

Maintenance and Operation

Pay close attention to the heating devices operating environment and always keep it clean so

that the heating devices can be used in a stabilized condition over a long period of time.

Operating the ibidi Heating Devices

Use the devices only in the intended application described in this manual.

CAUTION Unintended use which  is not in compliance with  this manual, can lead to a  failure

of the devices’ protective systems. This failure may cause personal injury or damage.

Cleaning the ibidi Heating Devices

Switch off the Temperature Controller and disconnect the heating devices from the controller.

To  clean the  devices  use only  a  dry cloth  or  a cloth  dampened with  water  (ultrapure)  or

common  lab  disinfection  solutions that  are  based  on  quaternary  ammonium  compounds

(e.g., Barrycidal 36 or Pharmacidal).

CAUTION The  use of  ethanol or  other types of  organic solvents  to clean  the devices  may

remove the devices’ paint.

Transporting the ibidi Heating Devices

First, switch off the Temperature Controller and then disconnect  the heating devices from the

controller. Carry the devices carefully and avoid mechanical shocks.

Repairing the ibidi Heating Devices

Do not try  to repair the devices  by yourself. Contact the  ibidi service to repair any  damaged

components.

CAUTION The  manual disassembling  of the devices  is not  allowed. Manual  disassembling

poses a  risk of personal  injury or  damage to the  devices. Contact ibidi  service personnel if

there is need to disassemble the devices.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

12

Waste Treatment

Waste treatment is your own  responsibility. You must dispose of the waste  using a company

that specializes in  waste recovery. Do not  dispose of the ibidi  heating devices in a  trash bin

or at a public waste disposable site. For detailed information, please contact the ibidi service.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

13

1   Working Principle

The stage  top incubator for  multi-well plates consists  of a heated  plate with a  heated glass

bottom, and also a  heated lid that is designed  to keep cells at 37°C  in multi-well plates. The

heated lid  prevents condensation effects inside  the entire incubation  system. Humidity, CO2

enriched,  and/or  O2   reduced/enriched  air  can  be  flushed   in,  which  makes  the  heated

chamber  a  fully controlled  incubator—right  on  the  microscope.  This  enables the  user  to

perform live cell imaging.

The glass  top of the  lid, the glass  bottom of the  plate and the  plate itself  are actively heat-

controlled. The  heated plate  has the  outer dimensions  of the  exchangeable centerpiece of

motorized stages that fit K-Frame Inserts (160 mm x 110 mm).

Figure 1: Schematic Drawing of the ibidi Heating devices

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

14

2   Overview

2.1   System Overview

Temperature Controller

Heating Devices

Optional: Computer

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

15

2.2   Components

ibidi Heated Glass Lid for Multi-Well

Plates (K-Frame)

ibidi Heated Plate with heated glass

bottom for Multi-Well Plates

(K-Frame)

Calibration Set

(1 temperature sensor + 1 perforated µ-Plate

96well)

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

16

2.3   Heated Lid

The heated lid (160 mm  x 108 mm) provides excellent optical quality, and  also allows for the

use of  all standard microscopy  techniques, including differential  interference contrast (DIC).

The upper  glass  part is  electrically heat  controlled. The  heated lid  fits  accurately onto  the

heated plate.

For gas  incubation, the  lid is  equipped with inlets  for (a)  the gas  flow and (b)  the humidity

sensor.

Due to the height of  the lid, we recommend using condensers with  a working distance of

30 mm.

>

2.4   Heated Plate with glass bottom

The heated  plate provides the base  for the heating  devices. Its format (160  mm x 110  mm)

equals exactly the exchangeable centerpiece of motorized stages in K-Frame Format.

The  glass  bottom and  the  plate  itself  are  both  actively heated.  The  heated  plate  is  the

counterpart to the heated lid.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

17

2.5   Sterilization and Cleaning

The  Heated Plate  and  the Heated  Lid  do not  have  to  be sterile,  because  the incubation

chamber has  no direct contact with  the cells and  the cell media. If  sterilization is necessary

for  some  reason,  we  recommend  using   isopropanol  (70%)  or  common  lab  disinfection

solutions based on quaternary ammonium compounds. (e.g. Barrycidal 36, Pharmacidal)

All  Heating Devices  can  be  cleaned. We  recommend  using  ultrapure  water for  cleaning.

Stains on  the  Heated Lid  or the  glass  bottom of  the Heated  Plate can  be  removed using

isopropanol (70%) or lens/eyeglass cleaning wipes.

Make sure  not  to apply  to much  pressure when  cleaning  the glass  bottom of  the plate  to

avoid damaging the glass.

2.6   Setting into Operation

Here is  the procedure  to start the  operation of  the heating  devices in combination  with the

ibidi Temperature Controller.

For the first power  up, we recommend disconnecting the  heating devices. By doing this, you

ensure that you do not start heating with the wrong temperature settings. It is now possible to

set the temperatures  for the individual  channels. Afterwards, you  can reconnect the heating

devices.

Since the controller is preconfigured,  we recommend connecting the heated lid  to Channel 1

the heated  plate to  Channel 2 and  the glass  bottom of  the heated plate  to Channel  3. For

subsequent power ups, you will not need to disconnect the heating devices.

Please  let the  temperature  of  the system  equilibrate  for  a

minimum of 30 minutes before you start your experiments.

Be  aware  that  when the  system  is  switched  on,  10V  AC

voltage is applied  to the underside of the  glass plate. Do not

touch the  underside  or contact  it with  anything conductive.

This   could  cause   a   short  circuit   that   may  destroy   the

controller and/or the lid.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

18

3   Temperature Calibration

The temperature at  the position of the  cells must be calibrated from  time to time, due  to the

following conditions:  room temperature, air conditioning,  heat loss via  microscope, chamber

type or objective lens, and humidification. A calibration is recommended before beginning the

first  experiment. The  heat  calibration should  be  done for  each  chamber type,  in  order  to

control temperatures with an absolute accuracy of less than 1°C.

For calibration, the ibidi  Temperature Controller has a plug, located on  the rear panel, for an

external temperature  sensor.  Before shipment,  all controllers  are  run through  an in-house

calibration  to  ensure  their  accuracy.   This  calibration  only  refers  to  the  actively  heated

devices, such as  the plate and the lid.  In most cases, the cell  temperature will lie somewhat

below the  temperature of  the heated plate.  To compensate for  this effect,  the heated  plate

can be calibrated  to the cell temperature that  is sensed by the  external temperature sensor,

which can be placed at the position of the cells.

The displayed  temperatures are valid  for the heating

devices  only. The  temperature near  the  cells might

be different (lower in most cases)!

3.1   Calibration Routine

The recommended  calibration routine is  fairly simple. Follow these  eight steps for  the initial

temperature calibration:

1)   Plug  the   external  temperature   sensor  into  the   rear  panel   of  the   Temperature

Controller.

2)   Fill the multi-well plate you want to use with medium and place the sensor as  close as

possible to  the  future position  of  the cells.  Use the  prepared  calibration set,  which

consists of a µ-Plate 96 well with a perforated lid, as shown in Figure 2 on page 20.

3)   Set all parameters  to create  conditions identical  to those  of the  actual experiments,

such as room temperature,  air conditioning, airflow, illumination, microscope  settings,

etc.

4)   Set the temperature of glass  bottom and heated plate to the  desired cell temperature

and  set  the temperature  of  the  lid  several  degrees higher  (recommended  values:

glass 37°, plate  37 °C, lid 42 °C).  You can do this directly  with the control buttons on

the controller, or with the TempControl software.

5)   Let the system equilibrate for 30 minutes.

6)   In most cases the temperature shown  by the external sensor will not  match the value

of the plate. In  that case increase or decrease the temperature  of the glass bottom of

the  plate  (recommended   in  0.5  °C   steps  for  small  deviations).   Let  the  system

equilibrate for another 5 minutes.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

19

Repeat this step until you reach the desired temperature at the external sensor.

7)   Note the adjusted set values in your documentation for future experiments.

8)   Check the temperature of the system from time to time.

It is recommended to perform this routine once every year.

It is  recommended to perform  this routine  when a change

of temperature conditions has occurred.

Figure 2: Sensor Placement for Temperature Calibration

3.2   Influence of Surrounding Temperature and Airflow

The  surrounding,   ambient  temperature   affects  the   temperature  inside   the  microscopy

chambers.  Devices, such  as  computers and  camera  controllers,  can significantly  heat  up

small rooms.  In this case,  we recommend  equilibrating the room  temperature to  the typical

experimental conditions at least 2-3 hours before calibrating.

Keep in mind that the temperature sensor can also measure room temperatures.

The airflow  can enhance the  effect of  temperature changes in  the vicinity of  the incubation

chamber.  In a  case where  the  airflow (e.g.,  air  conditioning, etc.)  cannot  be stopped,  we

recommend protecting the  microscope as much  as possible from any  strong airflow. Please

contact us for support.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

20

4   Troubleshooting

4.1   Focus Not Stable

Focus drift is  a disturbing effect,  especially during time-lapse experiments.  Focus stability is

mainly   influenced  by   mechanical   changes   and   temperature  variations.   Follow   these

recommendations to keep your cells in focus:











Switch  on all  components  (heating, gas  incubation,  computer, other  equipment)  at

least 60 minutes before starting the experiment.

After you put the sample onto  the microscope, wait 20 minutes before starting a time-

lapse experiment to achieve temperature and immersion oil equilibration. (*)

Keep the  room temperature  as stable as  possible. Air  conditioning should  either be

working continuously or switched off.

Do not change the temperature during the experiments. Avoid door/window  openings,

as this could rapidly change the temperature.

Eliminate  all   sources  of   mechanical   vibrations.  Use   a  damped   table  for   your

microscope.

(*) In  case the experiment  needs to be  started immediately,  either after placing  the sample

on  the  microscope,  or  after  closing the  lid,  we  recommend  controlling  the  focus  for  20

minutes.  In the  first minutes  after  starting the  experiment,  temperature equilibration  might

influence the focus/z-position of the cells.

4.2   Evaporation Too High

Depending on  the incubating  conditions, small volumes  might evaporate  quickly, especially

during long-term experiments.  We suggest using silicone  oil (e.g. Anti-Evaporation  Oil, ibidi,

50051)  to  decrease  evaporation.  If  you  have  an  actively   controlled  humidifying  device,

increase the set value of relative humidity.

4.2.1

Silicone Oil Procedure (ibidi’s Anti-Evaporation Oil)

Covering  the   medium  with   sterile  silicone   oil  prevents  all   evaporation  effects   and  is

compatible with cell culture. Please don’t use mineral  oil, as this would be harmful to the ibidi

µ-Slides.

Equilibrate oil and medium inside the incubator overnight. This step helps avoid the formation

of air bubbles, and  pre-warms the solutions to 37°C.  Afterwards, fill your plate with cells  and

medium. Cover  the medium’s  surface with an  appropriate amount  of silicone oil.  Don’t drip

the  oil directly  onto the  surface,  but let  it run  down  the edges  by  pressing the  pipette tip

directly on the upper side of the reservoir.

Please contact us at info@ibidi.com for further troubleshooting help.

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

© ibidi GmbH, Version 1.0, May 28th 2015

21

5   WEEE/RoHS Compliance Statement

EU Directives WEEE and RoHS

To Our Valued Customers:

ibidi GmbH  is committed to being  a good corporate  citizen. As part  of that commitment,  we

strive  to  maintain  an  environmentally conscious  manufacturing  operation.  The  European

Union (EU)  has enacted two  directives, the  first on product  recycling (Waste  Electrical and

Electronic  Equipment,  WEEE) and  the  second  on  limiting  the use  of  certain  substances

(Restriction on the use of Hazardous Substances, RoHS).

Two Categories  of products  covered by the  WEEE Directive  are currently exempt  from the

RoHS Directive  – Category  8, medical devices  (with the  exception of implanted  or infected

products) and Category 9, monitoring and control instruments.

All of our  products fall into  either Category 8 or  9, and are currently  exempt from the RoHS

Directive.  We will,  however,  continue  to monitor  the  application of  the  RoHS Directive  in

relation to our products and will comply with any changes, as they apply.

Recycling is offered for  our products that fall within  the scope of the WEEE  Directive. These

specific products,  available for sale  after August  13, 2005, will  be identified with  a “wheelie

bin” symbol.

When  you see  this symbol,  do  not dispose  of  this product  with  Municipal Waste.  Special

Collection/Disposal is required.

Please contact us for feedback or further information.

ibidi GmbH

Am Klopferspitz 19

82152 Planegg-Martinsried

Germany

www.ibidi.com

info@ibidi.de

Tel.: +49 89 520 4617 0

Fax: +49 89 520 4617 59

Heated Lid & Heated Plate for Multi-Well Plates (K-Frame–  160x110mm )

2

©ibidi GmbH, Version 1.0, May 28th 2015

22